## Results and Sample ID
Run Name:
Run Date:
Operator:
|  |  |  | Sample Info |  |  |  | Ammonium |  | Phosphate |  | Calibrant | Conc., mg/L
Row | SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | Peak# | Position | Identifier | Type | Raw Ht | ppm | Raw Ht | ppm | C1 | 0.002
|  |  |  | 1 | 1:1 | SYNC | SYNC | 0.037 | 0.0675 | 0.048 | 0.2005 | C2 | 0.005
|  |  |  | 2 | 1:2 | CO | Carry over | 0.001 | 0.0496 | 0 | 0.0002 | C3 | 0.01
|  |  |  | 3 | 1:3 | W | Wash | 0 | 0.0492 | 0 | -0.001 | C4 | 0.05
|  |  |  | 4 | 1:4 | C1 | Calibrant | 0.048 | 0.0715 | 0 | 0.0012 | C5 | 0.1
|  |  |  | 5 | 1:4 | C1 | Calibrant | 0.049 | 0.0706 | 0 | 0.0013 | C6 | 0.2
|  |  |  | 6 | 1:4 | C1 | Calibrant | 0.041 | 0.0654 | 0 | 0.0014
|  |  |  | 7 | 1:5 | C2 | Calibrant | 0.001 | 0.0439 | 0.001 | 0.0045
|  |  |  | 8 | 1:5 | C2 | Calibrant | 0.003 | 0.0438 | 0.001 | 0.0049
|  |  |  | 9 | 1:5 | C2 | Calibrant | 0.001 | 0.0414 | 0.001 | 0.005
|  |  |  | 10 | 1:6 | C3 | Calibrant | 0.001 | 0.0398 | 0.002 | 0.0095
|  |  |  | 11 | 1:6 | C3 | Calibrant | 0.037 | 0.0564 | 0.002 | 0.0096
|  |  |  | 12 | 1:6 | C3 | Calibrant | 0.002 | 0.0366 | 0.002 | 0.01
|  |  |  | 13 | 1:7 | C4 | Calibrant | 0.002 | 0.0354 | 0.012 | 0.051
|  |  |  | 14 | 1:7 | C4 | Calibrant | 0.009 | 0.0373 | 0.012 | 0.0509
|  |  |  | 15 | 1:7 | C4 | Calibrant | 0.009 | 0.0357 | 0.012 | 0.0513
|  |  |  | 16 | 1:8 | C5 | Calibrant | 0.009 | 0.0341 | 0.024 | 0.1008
|  |  |  | 17 | 1:8 | C5 | Calibrant | 0.02 | 0.0377 | 0.024 | 0.1009
|  |  |  | 18 | 1:8 | C5 | Calibrant | 0.021 | 0.0364 | 0.024 | 0.1006
|  |  |  | 19 | 1:9 | C6 | Calibrant | 0.021 | 0.0349 | 0.048 | 0.2003
|  |  |  | 20 | 1:9 | C6 | Calibrant | 1.192 | 0.6078 | 0.047 | 0.1984
|  |  |  | 21 | 1:9 | C6 | Calibrant | 0.1 | 0.0589 | 0.047 | 0.1994
|  |  |  | 22 | 1:3 | W | Wash | 0.064 | 0.0492 | -0.001 | -0.001
|  |  |  | 23 | 1:11 | CC1 | Unknown | 0.082 | 0.0582 | 0.023 | 0.1007
| 2018-07-18 | 13 | a | 24 | 2:1 | 37 | Unknown | 0.065 | 0.0495 | 0.001 | 0.0057
| 2018-07-18 | 13 | b | 25 | 2:2 | 38 | Unknown | 0.065 | 0.0497 | 0 | 0.0039
| 2018-07-18 | 13 | c | 26 | 2:3 | 39 | Unknown | 0.064 | 0.0495 | 0 | 0.0048
| 2018-08-02 | 1 | a | 27 | 2:4 | 40 | Unknown | 0.065 | 0.0499 | 0.004 | 0.0203
| 2018-08-02 | 1 | b | 28 | 2:5 | 41 | Unknown | 0.065 | 0.05 | 0.003 | 0.0176
| 2018-08-02 | 1 | c | 29 | 2:6 | 42 | Unknown | 0.065 | 0.0498 | 0.002 | 0.0123
| 2018-08-02 | 2 | a | 30 | 2:7 | 43 | Unknown | 0.065 | 0.0498 | 0.004 | 0.0192
| 2018-08-02 | 2 | b | 31 | 2:8 | 44 | Unknown | 0.065 | 0.0497 | 0.003 | 0.0175
| 2018-08-02 | 2 | c | 32 | 2:9 | 45 | Unknown | 0.065 | 0.0497 | 0.004 | 0.0185
| 2018-08-02 | 3 | a | 35 | 2:10 | 46 | Unknown | 0.065 | 0.0496 | 0.004 | 0.0186
| 2018-08-02 | 3 | b | 36 | 2:11 | 47 | Unknown | 0.065 | 0.0497 | 0.004 | 0.021
| 2018-08-02 | 3 | c | 37 | 2:12 | 48 | Unknown | 0.065 | 0.0497 | 0.005 | 0.0229
| 2018-08-02 | 4 | a | 38 | 2:13 | 49 | Unknown | 0.065 | 0.0498 | 0.005 | 0.0244
| 2018-08-02 | 4 | b | 39 | 2:14 | 50 | Unknown | 0.064 | 0.0496 | 0.002 | 0.0128
| 2018-08-02 | 4 | c | 40 | 2:15 | 51 | Unknown | 0.064 | 0.0496 | 0.003 | 0.0142
| 2018-08-02 | 5 | a | 41 | 2:16 | 52 | Unknown | 0.065 | 0.0497 | 0.002 | 0.0131
| 2018-08-02 | 5 | b | 42 | 2:17 | 53 | Unknown | 0.064 | 0.0497 | 0.002 | 0.0133
| 2018-08-02 | 5 | c | 43 | 2:18 | 54 | Unknown | 0.064 | 0.0496 | 0.001 | 0.0097
| 2018-08-02 | 6 | a | 45 | 2:19 | 55 | Unknown | 0.064 | 0.0497 | 0.001 | 0.0084
| 2018-08-02 | 6 | b | 46 | 2:20 | 56 | Unknown | 0.065 | 0.0499 | 0.001 | 0.009
| 2018-08-02 | 6 | c | 47 | 2:21 | 57 | Unknown | 0.065 | 0.0498 | 0.001 | 0.0088
| 2018-08-02 | 7 | a | 48 | 2:22 | 58 | Unknown | 0.065 | 0.0498 | 0.001 | 0.0078
| 2018-08-02 | 7 | b | 49 | 2:23 | 59 | Unknown | 0.065 | 0.0498 | 0.001 | 0.0075
| 2018-08-02 | 7 | c | 50 | 2:24 | 60 | Unknown | 0.065 | 0.05 | 0.001 | 0.0092
| 2018-08-02 | 8 | a | 51 | 2:25 | 61 | Unknown | 0.107 | 0.0707 | 0.001 | 0.0083
| 2018-08-02 | 8 | b | 52 | 2:26 | 62 | Unknown | 0.064 | 0.0493 | 0.001 | 0.0087
| 2018-08-02 | 8 | c | 53 | 2:27 | 63 | Unknown | 0.064 | 0.0495 | 0.001 | 0.0066
| 2018-08-02 | 9 | a | 54 | 2:28 | 64 | Unknown | 0.064 | 0.0494 | 0.001 | 0.01
| 2018-08-02 | 9 | b | 56 | 2:29 | 65 | Unknown | 0.064 | 0.0497 | 0.002 | 0.0133
| 2018-08-02 | 9 | c | 57 | 2:30 | 66 | Unknown | 0.064 | 0.0496 | 0.002 | 0.0105
| 2018-08-02 | 10 | a | 58 | 2:31 | 67 | Unknown | 0.064 | 0.0495 | 0.002 | 0.0143
| 2018-08-02 | 10 | b | 59 | 2:32 | 68 | Unknown | 0.064 | 0.0495 | 0.002 | 0.0123
| 2018-08-02 | 10 | c | 60 | 2:33 | 69 | Unknown | 0.064 | 0.0495 | 0.002 | 0.0119
| 2018-08-02 | 11 | a | 61 | 2:34 | 70 | Unknown | 0.064 | 0.0495 | 0.003 | 0.0151
| 2018-08-02 | 11 | b | 62 | 2:35 | 71 | Unknown | 0.064 | 0.0496 | 0.002 | 0.0138
| 2018-08-02 | 11 | c | 63 | 2:36 | 72 | Unknown | 0.064 | 0.0497 | 0.003 | 0.0168
| 2018-08-02 | 12 | a | 64 | 2:37 | 73 | Unknown | 0.064 | 0.0497 | 0.002 | 0.0121
| 2018-08-02 | 12 | b | 65 | 2:38 | 74 | Unknown | 0.064 | 0.0497 | 0.002 | 0.0134
| 2018-08-02 | 12 | c | 67 | 2:39 | 75 | Unknown | 0.064 | 0.0496 | 0.001 | 0.0094
| 2018-08-02 | 13 | a | 69 | 2:40 | 76 | Unknown | 0.064 | 0.0495 | 0 | 0.0041
| 2018-08-02 | 13 | b | 70 | 2:41 | 77 | Unknown | 0.064 | 0.0496 | 0 | 0.0057
| 2018-08-02 | 13 | c | 71 | 2:42 | 78 | Unknown | 0.064 | 0.0496 | 0 | 0.005
| 2018-08-16 | 1 | a | 72 | 2:43 | 79 | Unknown | 0.064 | 0.0496 | 0.004 | 0.0206
| 2018-08-16 | 1 | b | 73 | 2:44 | 80 | Unknown | 0.064 | 0.0495 | 0.003 | 0.0173
| 2018-08-16 | 1 | c | 74 | 2:45 | 81 | Unknown | 0.064 | 0.0495 | 0.003 | 0.0167
| 2018-08-16 | 2 | a | 75 | 2:46 | 82 | Unknown | 0.065 | 0.0498 | 0.003 | 0.0167
| 2018-08-16 | 2 | b | 76 | 2:47 | 83 | Unknown | 0.065 | 0.0499 | 0.003 | 0.016
| 2018-08-16 | 2 | c | 78 | 2:48 | 84 | Unknown | 0.065 | 0.0499 | 0.004 | 0.0189
| 2018-08-16 | 3 | a | 79 | 2:49 | 85 | Unknown | 0.065 | 0.05 | 0.004 | 0.0194
| 2018-08-16 | 3 | b | 80 | 2:50 | 86 | Unknown | 0.065 | 0.0499 | 0.003 | 0.0175
| 2018-08-16 | 3 | c | 81 | 2:51 | 87 | Unknown | 0.065 | 0.0499 | 0.003 | 0.0169
| 2018-08-16 | 4 | a | 82 | 2:52 | 88 | Unknown | 0.065 | 0.0499 | 0.002 | 0.0116
| 2018-08-16 | 4 | b | 83 | 2:53 | 89 | Unknown | 0.064 | 0.0497 | 0.002 | 0.0121
| 2018-08-16 | 4 | c | 84 | 2:54 | 90 | Unknown | 0.065 | 0.0499 | 0.003 | 0.0147
| 2018-08-16 | 5 | a | 85 | 2:55 | 91 | Unknown | 0.064 | 0.0494 | 0.002 | 0.0093
| 2018-08-16 | 5 | b | 86 | 2:56 | 92 | Unknown | 0.064 | 0.0495 | 0.001 | 0.0079
| 2018-08-16 | 5 | c | 87 | 2:57 | 93 | Unknown | 0.064 | 0.0496 | 0.001 | 0.0068
| 2018-08-16 | 6 | a | 89 | 2:58 | 94 | Unknown | 0.064 | 0.0496 | 0 | 0.0044
| 2018-08-16 | 6 | b | 90 | 2:59 | 95 | Unknown | 0.064 | 0.0496 | 0.022 | 0.0931
| 2018-08-16 | 6 | c | 91 | 2:60 | 96 | Unknown | 0.064 | 0.0497 | 0.004 | 0.0171
| 2018-08-16 | 7 | a | 92 | 3:1 | 97 | Unknown | 0.064 | 0.0497 | 0 | 0.0036
| 2018-08-16 | 7 | b | 93 | 3:2 | 98 | Unknown | 0.064 | 0.0497 | 0 | 0.0018
| 2018-08-16 | 7 | c | 94 | 3:3 | 99 | Unknown | 0.064 | 0.0496 | 0 | 0.0033
| 2018-08-16 | 8 | a | 95 | 3:4 | 100 | Unknown | 0.064 | 0.0494 | 0 | 0.0022
| 2018-08-16 | 8 | b | 96 | 3:5 | 101 | Unknown | 0.064 | 0.0495 | 0.001 | 0.0075
| 2018-08-16 | 8 | c | 97 | 3:6 | 102 | Unknown | 0.064 | 0.0495 | 0 | 0.0042
| 2018-08-16 | 9 | a | 98 | 3:7 | 103 | Unknown | 0.064 | 0.0495 | 0.002 | 0.0112
| 2018-08-16 | 9 | b | 100 | 3:8 | 104 | Unknown | 0.065 | 0.0501 | 0.002 | 0.0093
| 2018-08-16 | 9 | c | 101 | 3:9 | 105 | Unknown | 0.064 | 0.0496 | 0.001 | 0.0085
| 2018-08-16 | 10 | a | 102 | 3:10 | 106 | Unknown | 0.064 | 0.0495 | 0.002 | 0.0126
| 2018-08-16 | 10 | b | 103 | 3:11 | 107 | Unknown | 0.064 | 0.0494 | 0.002 | 0.0124
| 2018-08-16 | 10 | c | 104 | 3:12 | 108 | Unknown | 0.064 | 0.0495 | 0.001 | 0.0087
| 2018-08-16 | 11 | a | 105 | 3:13 | 109 | Unknown | 0.064 | 0.0498 | 0.001 | 0.0063
| 2018-08-16 | 11 | b | 106 | 3:14 | 110 | Unknown | 0.064 | 0.0496 | 0.001 | 0.0084
| 2018-08-16 | 11 | c | 107 | 3:15 | 111 | Unknown | 0.064 | 0.0494 | 0.001 | 0.0072
| 2018-08-16 | 12 | a | 108 | 3:16 | 112 | Unknown | 0.065 | 0.0498 | 0.001 | 0.0047
| 2018-08-16 | 12 | b | 109 | 3:17 | 113 | Unknown | 0.064 | 0.0496 | 0.001 | 0.005
| 2018-08-16 | 12 | c | 111 | 3:18 | 114 | Unknown | 0.064 | 0.0496 | 0 | 0.0031
| 2018-08-16 | 13 | a | 112 | 3:19 | 115 | Unknown | 0.064 | 0.0497 | 0 | 0.0012
| 2018-08-16 | 13 | b | 113 | 3:20 | 116 | Unknown | 0.064 | 0.0496 | 0 | 0.0007
| 2018-08-16 | 13 | c | 114 | 3:21 | 117 | Unknown | 0.064 | 0.0496 | 0 | 0.0013
| 2018-08-31 | 1 | a | 115 | 3:22 | 118 | Unknown | 0.064 | 0.0496 | 0.002 | 0.0125
| 2018-08-31 | 1 | b | 116 | 3:23 | 119 | Unknown | 0.064 | 0.0496 | 0.003 | 0.0145
| 2018-08-31 | 1 | c | 117 | 3:24 | 120 | Unknown | 0.064 | 0.0496 | 0.003 | 0.0132
| 2018-08-31 | 2 | a | 118 | 3:25 | 121 | Unknown | 0.065 | 0.0498 | 0.003 | 0.0133
| 2018-08-31 | 2 | b | 119 | 3:26 | 122 | Unknown | 0.065 | 0.0498 | 0.07 | 0.2928
| 2018-08-31 | 2 | c | 120 | 3:27 | 123 | Unknown | 0.064 | 0.0496 | 0.003 | 0.0131
| 2018-08-31 | 3 | a | 122 | 3:28 | 124 | Unknown | 0.099 | 0.0667 | 0.003 | 0.0126
| 2018-08-31 | 3 | b | 123 | 3:29 | 125 | Unknown | 0.064 | 0.0492 | 0.002 | 0.0105
| 2018-08-31 | 3 | c | 124 | 3:30 | 126 | Unknown | 0.064 | 0.0496 | 0.049 | 0.205
| 2018-08-31 | 4 | a | 125 | 3:31 | 127 | Unknown | 0.064 | 0.0497 | 0.003 | 0.0099
| 2018-08-31 | 4 | b | 126 | 3:32 | 128 | Unknown | 0.065 | 0.05 | 0.069 | 0.2868
| 2018-08-31 | 4 | c | 127 | 3:33 | 129 | Unknown | 0.064 | 0.0496 | 0.004 | 0.0132
| 2018-08-31 | 5 | a | 128 | 3:34 | 130 | Unknown | 0.064 | 0.0497 | 0.003 | 0.0072
| 2018-08-31 | 5 | b | 129 | 3:35 | 131 | Unknown | 0.064 | 0.0496 | 0.163 | 0.6751
| 2018-08-31 | 5 | c | 130 | 3:36 | 132 | Unknown | 0.064 | 0.0497 | 0.001 | -0.0028
| 2018-08-31 | 6 | a | 131 | 3:37 | 133 | Unknown | 0.064 | 0.0496 | 0.443 | 1.8361
| 2018-08-31 | 6 | b | 134 | 3:38 | 134 | Unknown | 0.065 | 0.0496 | 0.14 | 0.5783
| 2018-08-31 | 6 | c | 135 | 3:39 | 135 | Unknown | 0.064 | 0.0496 | 0.038 | 0.1494
| 2018-08-31 | 7 | a | 136 | 3:40 | 136 | Unknown | 0.064 | 0.0496 | 0.003 | 0.0072
| 2018-08-31 | 7 | b | 137 | 3:41 | 137 | Unknown | 0.065 | 0.0501 | 0.002 | 0.0074
| 2018-08-31 | 7 | c | 138 | 3:42 | 138 | Unknown | 0.064 | 0.0496 | 0.001 | 0.003
| 2018-08-31 | 8 | a | 139 | 3:43 | 139 | Unknown | 0.064 | 0.0496 | 0.035 | 0.1434
| 2018-08-31 | 8 | b | 140 | 3:44 | 140 | Unknown | 0.065 | 0.0497 | 0.028 | 0.1137
| 2018-08-31 | 8 | c | 141 | 3:45 | 141 | Unknown | 0.064 | 0.0495 | 0.001 | 0.0065
| 2018-08-31 | 9 | a | 142 | 3:46 | 142 | Unknown | 0.064 | 0.0496 | 0.023 | 0.0969
| 2018-08-31 | 9 | b | 144 | 3:47 | 143 | Unknown | 0.064 | 0.0493 | 0.12 | 0.5024
| 2018-08-31 | 9 | c | 145 | 3:48 | 144 | Unknown | 0.064 | 0.0494 | 0.145 | 0.6007
| 2018-08-31 | 10 | a | 146 | 3:49 | 145 | Unknown | 0.064 | 0.0495 | 0.157 | 0.6508
| 2018-08-31 | 10 | b | 147 | 3:50 | 146 | Unknown | 0.064 | 0.0494 | 0.036 | 0.1466
| 2018-08-31 | 10 | c | 148 | 3:51 | 147 | Unknown | 0.064 | 0.0495 | 0.149 | 0.6197
| 2018-08-31 | 11 | a | 149 | 3:52 | 148 | Unknown | 0.064 | 0.0494 | 0.304 | 1.2595
| 2018-08-31 | 11 | b | 150 | 3:53 | 149 | Unknown | 0.064 | 0.0494 | 0.101 | 0.4113
| 2018-08-31 | 11 | c | 151 | 3:54 | 150 | Unknown | 0.064 | 0.0494 | 0.102 | 0.4199
| 2018-08-31 | 12 | a | 152 | 3:55 | 151 | Unknown | 0.064 | 0.0496 | 0.195 | 0.8072
| 2018-08-31 | 12 | b | 153 | 3:56 | 152 | Unknown | 0.064 | 0.0496 | 0.002 | 0.0005
| 2018-08-31 | 12 | c | 155 | 3:57 | 153 | Unknown | 0.064 | 0.0496 | 0 | -0.0019
| 2018-08-31 | 13 | a | 156 | 3:58 | 154 | Unknown | 0.064 | 0.0496 | -0.001 | -0.0059
| 2018-08-31 | 13 | b | 157 | 3:59 | 155 | Unknown | 0.064 | 0.0496 | 0 | -0.0032
| 2018-08-31 | 13 | c | 158 | 3:60 | 156 | Unknown | 0.076 | 0.0551 | ??? | ???

## Std Curve Construction
1. Copy in Calibrant Information | [5 empty] | NH3 SECTION | [6 empty] | PO4 SECTION
[18 empty] | 2. Run Linear Regression | [8 empty] | 4. If quadratic exceeds 10%, run split-curve procedure.
Use X=Peak height; Y=Conc)
| Calibrant |  |  | Linear Fit |  | Quadratic |  | Split Curve |  |  | Linear Fit |  | Quadratic |  | Split Curve
Calibrant | Conc. | NH3 Calibrant | PO4 Calibrant | Std Conc |  | Std Conc |  | Std Conc |  |  | Std Conc |  | Std Conc |  | Std Conc | [11 empty] | Plot low range calibrant data: |  |  | Enter coefficients below graph.
Id | (mg/L N) | peak height (x) | peak height (x) | (y-hat) | % error | (y-hat) | % error | (y-hat) | % error |  | (y-hat) | % error | (y-hat) | % error | (y-hat) | % error |  |  | All data - Linear Fit
C1 | 0.002 | 0.048 | 0 | 0.055977 | 2698.8704 | 0.52758 | 26279.008 | 0.6094 | 30370 |  | 0.000962 | -51.9 | 0.0016 | -20 | 0.001 | -50 |  | SUMMARY OUTPUT
C1 | 0.002 | 0.049 | 0 | 0.05611 | 2705.524 | 0.54263 | 10752.6518 | 0.6271 | 12442 |  | 0.000962 | -51.9 | 0.0016 | -20 | 0.001 | -80
C1 | 0.002 | 0.041 | 0 | 0.055046 | 2652.2955 | 0.42669 | 4166.9259 | 0.4911 | 4811 |  | 0.000962 | -51.9 | 0.0016 | -20 | 0.001 | -90 |  | Regression Statistics
C2 | 0.005 | 0.001 | 0.001 | 0.049723 | 894.4614 | 0.00057339 | -98.8532 | 0.0031 | -93.8 |  | 0.0051461 | 2.9215 | 0.0056037 | 12.0733 | 0.002 | -96 |  | Multiple R | 0.99984
C2 | 0.005 | 0.003 | 0.001 | 0.049989 | 899.7843 | 0.0158 | -84.1999 | 0.0199 | -80.1 |  | 0.0051461 | 2.9215 | 0.0056037 | 12.0733 | 0.002 | -98 |  | R Square | 0.99968
C2 | 0.005 | 0.001 | 0.001 | 0.049723 | 894.4614 | 0.00057339 | -99.7133 | 0.0031 | -98.45 |  | 0.0051461 | 2.9215 | 0.0056037 | 12.0733 | 0.002 | -99 |  | Adjusted R Square | 0.99966
C3 | 0.01 | 0.001 | 0.002 | 0.049723 | 397.2307 | 0.00057339 | #REF! | 0.0031 | #REF! |  | 0.0093302 | -6.6985 | 0.0096155 | -3.8455 | 0.005 | #REF! |  | Standard Error | 0.0013484
C3 | 0.01 | 0.037 | 0.002 | 0.054514 | 445.1363 | 0.37256 | 3625.6211 | 0.4279 | 4179 |  | 0.0093302 | -6.6985 | 0.0096155 | -3.8455 | 0.005 | -50 |  | Observations | 18
C3 | 0.01 | 0.002 | 0.002 | 0.049856 | 398.5614 | 0.0081068 | -18.9324 | 0.0114 | 14 | Split | 0.0093302 | -6.6985 | 0.0096155 | -3.8455 | 0.005 | -50 | Split
C4 | 0.05 | 0.002 | 0.012 | 0.049856 | -0.28772 | 0.0081068 | #REF! | 0.0087934 | #REF! |  | 0.051171 | 2.3418 | 0.05018 | 0.36052 | 0.042063 | #REF! |  | ANOVA
C4 | 0.05 | 0.009 | 0.012 | 0.050788 | 1.5753 | 0.06532 | 30.6396 | 0.085502 | 71.0037 |  | 0.051171 | 2.3418 | 0.05018 | 0.36052 | 0.042063 | -15.8732 |  |  | df | SS | MS | F | Significance F
C4 | 0.05 | 0.009 | 0.012 | 0.050788 | 1.5753 | 0.06532 | 30.6396 | 0.085502 | 71.0037 |  | 0.051171 | 2.3418 | 0.05018 | 0.36052 | 0.042063 | -15.8732 |  | Regression | 1 | 0.090513 | 0.090513 | 49785.8232 | 2.2292e-29
C5 | 0.1 | 0.009 | 0.024 | 0.050788 | -49.2124 | 0.06532 | #REF! | 0.085502 | #REF! |  | 0.10138 | 1.3798 | 0.099931 | -0.06937 | 0.094336 | #REF! |  | Residual | 16 | 2.9089e-05 | 1.8181e-06
C5 | 0.1 | 0.02 | 0.024 | 0.052251 | -47.7486 | 0.17106 | 71.064 | 0.247 | 147 |  | 0.10138 | 1.3798 | 0.099931 | -0.06937 | 0.094336 | -5.6641 |  | Total | 17 | 0.090543
C5 | 0.1 | 0.021 | 0.024 | 0.052384 | -47.6155 | 0.18164 | #REF! | 0.26416 | #REF! |  | 0.10138 | 1.3798 | 0.099931 | -0.06937 | 0.094336 | #REF!
C6 | 0.2 | 0.021 | 0.048 | 0.088828 | -55.586 | 0.18164 | -9.1815 | 0.2071 | 3.55 |  | 0.2018 | 0.89882 | 0.20294 | 1.4709 | 0.6094 | 204.7 |  |  | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
C6 | 0.2 | 1.192 | 0.047 | 4.9884 | 2394.1903 | 122.3418 | 61070.9221 | 151.6174 | 75708.7 |  | 0.19761 | -1.1932 | 0.19856 | -0.72193 | 0.5919 | 195.95 |  | Intercept | 0.00096247 | 0.0004169 | 2.3086 | 0.034652 | 7.8678e-05 | 0.0018463 | 7.8678e-05 | 0.0018463 | NH3 Low curve coefficients:
C6 | 0.2 | 0.1 | 0.047 | 0.41937 | 109.685 | 1.5224 | #REF! | 1.795 | #REF! |  | 0.19761 | -1.1932 | 0.19856 | -0.72193 | 0.5919 | #REF! |  | X Variable 1 | 4.1841 | 0.018752 | 223.1274 | 2.2292e-29 | 4.1443 | 4.2238 | 4.1443 | 4.2238 |  | intcpt | X coef. | X^2 coef.
[28 empty] | -0.005 | 8 | 100
PO4 Low curve coefficients:
[18 empty] | Record Int and Slope here: | [9 empty] | intcpt | X coef. | X^2 coef.
[18 empty] | NH3 Int. | 0.04959 |  | PO4 Int. | 0.000962 | [5 empty] | 0.001 | 1e-13 | 1000
[18 empty] | NH3 Slope | 0.13307 |  | PO4 Slope | 4.1841
[27 empty] | Plot High Range calibrant data. |  |  | Enter coefficients below graph.
3. If errors exceed 10% from above, run quadratic fit.
(Use X=Peak Ht.; Y=Conc)
Enter equation parameters below graph.
NH3High curve coefficients:
[28 empty] | intcpt | X coef. | X^2 coef.
[28 empty] | -0.0094 | 8.683 | 206.85
NH3 Quadratic curve coefficients
[21 empty] | Intercept | X coef. | X^2 coef. |  |  |  | PO4High curve coefficients:
[21 empty] | -0.0068 | 7.2934 | 79.99 |  |  |  |  | intcpt | X coef. | X^2 coef.
[28 empty] | 0.0007 | 2.9924 | 37.879
PO4 Quadratic curve coefficients
[21 empty] | Intercept | X coef. | X^2 coef.
[21 empty] | 0.0016 | 3.9996 | 4.0629

## Apply Std Curve
|  |  | Sample Info |  |  |  | FASPac NH3 |  | FASPac PO4 |  | NH3 curve fit | PO4 curve fit | NH3 |  |  |  | PO4
SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | Peak# | Position | Identifier | Type | Raw Ht | mg/L | Raw Ht | mg/L | Results mg/L | Results mg/L | Curve-fitting parameters used: |  |  |  | Curve-fitting parameters used:
2018-07-18 | 13 | a | 24 | 2:1 | 37 | Unknown | 0.065 | 0.0495 | 0.001 | 0.0057 | 0.41666 | 0.0056037 | Copy from Std Curve Construction page. |  |  |  | Copy from Std Curve Construction page.
2018-07-18 | 13 | b | 25 | 2:2 | 38 | Unknown | 0.065 | 0.0497 | 0 | 0.0039 | 0.41666 | 0.0016
2018-07-18 | 13 | c | 26 | 2:3 | 39 | Unknown | 0.064 | 0.0495 | 0 | 0.0048 | 0.41016 | 0.0016 | Use these for linear or quadratic fit. |  |  |  | Use these for linear or quadratic fit.
2018-08-02 | 1 | a | 27 | 2:4 | 40 | Unknown | 0.065 | 0.0499 | 0.004 | 0.0203 | 0.41666 | 0.017663 | If linear, enter 0 for X^2 coefficient. |  |  |  | If linear, enter 0 for X^2 coefficient.
2018-08-02 | 1 | b | 28 | 2:5 | 41 | Unknown | 0.065 | 0.05 | 0.003 | 0.0176 | 0.41666 | 0.013635
2018-08-02 | 1 | c | 29 | 2:6 | 42 | Unknown | 0.065 | 0.0498 | 0.002 | 0.0123 | 0.41666 | 0.0096155 | Intercept: | X Coef. | X^2 Coef: |  | Intercept: | X Coef. | X^2 Coef:
2018-08-02 | 2 | a | 30 | 2:7 | 43 | Unknown | 0.065 | 0.0498 | 0.004 | 0.0192 | 0.41666 | 0.017663 | -0.00555 | 6.4955 | 0 |  | 0.0016 | 3.9996 | 4.0629
2018-08-02 | 2 | b | 31 | 2:8 | 44 | Unknown | 0.065 | 0.0497 | 0.003 | 0.0175 | 0.41666 | 0.013635
2018-08-02 | 2 | c | 32 | 2:9 | 45 | Unknown | 0.065 | 0.0497 | 0.004 | 0.0185 | 0.41666 | 0.017663
2018-08-02 | 3 | a | 35 | 2:10 | 46 | Unknown | 0.065 | 0.0496 | 0.004 | 0.0186 | 0.41666 | 0.017663 | Use these for split curve: |  |  |  | Use these for split curve:
2018-08-02 | 3 | b | 36 | 2:11 | 47 | Unknown | 0.065 | 0.0497 | 0.004 | 0.021 | 0.41666 | 0.017663 | Low Range: |  |  |  | Low Range:
2018-08-02 | 3 | c | 37 | 2:12 | 48 | Unknown | 0.065 | 0.0497 | 0.005 | 0.0229 | 0.41666 | 0.0217 | Intercept: | X Coef. | X^2 Coef: |  | Intercept: | X Coef. | X^2 Coef:
2018-08-02 | 4 | a | 38 | 2:13 | 49 | Unknown | 0.065 | 0.0498 | 0.005 | 0.0244 | 0.41666 | 0.0217
2018-08-02 | 4 | b | 39 | 2:14 | 50 | Unknown | 0.064 | 0.0496 | 0.002 | 0.0128 | 0.41016 | 0.0096155
2018-08-02 | 4 | c | 40 | 2:15 | 51 | Unknown | 0.064 | 0.0496 | 0.003 | 0.0142 | 0.41016 | 0.013635
2018-08-02 | 5 | a | 41 | 2:16 | 52 | Unknown | 0.065 | 0.0497 | 0.002 | 0.0131 | 0.41666 | 0.0096155 | High Range: |  |  |  | High Range:
2018-08-02 | 5 | b | 42 | 2:17 | 53 | Unknown | 0.064 | 0.0497 | 0.002 | 0.0133 | 0.41016 | 0.0096155 | Intercept: | X Coef. | X^2 Coef: |  | Intercept: | X Coef. | X^2 Coef:
2018-08-02 | 5 | c | 43 | 2:18 | 54 | Unknown | 0.064 | 0.0496 | 0.001 | 0.0097 | 0.41016 | 0.0056037
2018-08-02 | 6 | a | 45 | 2:19 | 55 | Unknown | 0.064 | 0.0497 | 0.001 | 0.0084 | 0.41016 | 0.0056037
2018-08-02 | 6 | b | 46 | 2:20 | 56 | Unknown | 0.065 | 0.0499 | 0.001 | 0.009 | 0.41666 | 0.0056037 | Procedure for split curve: |  |  |  | Procedure for split curve:
2018-08-02 | 6 | c | 47 | 2:21 | 57 | Unknown | 0.065 | 0.0498 | 0.001 | 0.0088 | 0.41666 | 0.0056037 | 1. Sort data on Row I: FASpac results, low to high. |  |  |  | 1. Sort data on Row I: FASpac results, low to high.
2018-08-02 | 7 | a | 48 | 2:22 | 58 | Unknown | 0.065 | 0.0498 | 0.001 | 0.0078 | 0.41666 | 0.0056037 | 2. Apply low range equation. |  |  |  | 2. Apply low range equation.
2018-08-02 | 7 | b | 49 | 2:23 | 59 | Unknown | 0.065 | 0.0498 | 0.001 | 0.0075 | 0.41666 | 0.0056037 | 3. Apply high range equation. |  |  |  | 3. Apply high range equation.
2018-08-02 | 7 | c | 50 | 2:24 | 60 | Unknown | 0.065 | 0.05 | 0.001 | 0.0092 | 0.41666 | 0.0056037 | 4. Note here peak height where split occurs: |  |  | 0.05 | 4. Note here peak height where split occurs: |  |  | 0.05
2018-08-02 | 8 | a | 51 | 2:25 | 61 | Unknown | 0.107 | 0.0707 | 0.001 | 0.0083 | 0.68947 | 0.0056037 | 4. Sort back based on Row D: Peak Number |  |  |  | 4. Sort back based on Row D: Peak Number
2018-08-02 | 8 | b | 52 | 2:26 | 62 | Unknown | 0.064 | 0.0493 | 0.001 | 0.0087 | 0.41016 | 0.0056037
2018-08-02 | 8 | c | 53 | 2:27 | 63 | Unknown | 0.064 | 0.0495 | 0.001 | 0.0066 | 0.41016 | 0.0056037
2018-08-02 | 9 | a | 54 | 2:28 | 64 | Unknown | 0.064 | 0.0494 | 0.001 | 0.01 | 0.41016 | 0.0056037
2018-08-02 | 9 | b | 56 | 2:29 | 65 | Unknown | 0.064 | 0.0497 | 0.002 | 0.0133 | 0.41016 | 0.0096155
2018-08-02 | 9 | c | 57 | 2:30 | 66 | Unknown | 0.064 | 0.0496 | 0.002 | 0.0105 | 0.41016 | 0.0096155
2018-08-02 | 10 | a | 58 | 2:31 | 67 | Unknown | 0.064 | 0.0495 | 0.002 | 0.0143 | 0.41016 | 0.0096155
2018-08-02 | 10 | b | 59 | 2:32 | 68 | Unknown | 0.064 | 0.0495 | 0.002 | 0.0123 | 0.41016 | 0.0096155
2018-08-02 | 10 | c | 60 | 2:33 | 69 | Unknown | 0.064 | 0.0495 | 0.002 | 0.0119 | 0.41016 | 0.0096155
2018-08-02 | 11 | a | 61 | 2:34 | 70 | Unknown | 0.064 | 0.0495 | 0.003 | 0.0151 | 0.41016 | 0.013635
2018-08-02 | 11 | b | 62 | 2:35 | 71 | Unknown | 0.064 | 0.0496 | 0.002 | 0.0138 | 0.41016 | 0.0096155
2018-08-02 | 11 | c | 63 | 2:36 | 72 | Unknown | 0.064 | 0.0497 | 0.003 | 0.0168 | 0.41016 | 0.013635
2018-08-02 | 12 | a | 64 | 2:37 | 73 | Unknown | 0.064 | 0.0497 | 0.002 | 0.0121 | 0.41016 | 0.0096155
2018-08-02 | 12 | b | 65 | 2:38 | 74 | Unknown | 0.064 | 0.0497 | 0.002 | 0.0134 | 0.41016 | 0.0096155
2018-08-02 | 12 | c | 67 | 2:39 | 75 | Unknown | 0.064 | 0.0496 | 0.001 | 0.0094 | 0.41016 | 0.0056037
2018-08-02 | 13 | a | 69 | 2:40 | 76 | Unknown | 0.064 | 0.0495 | 0 | 0.0041 | 0.41016 | 0.0016
2018-08-02 | 13 | b | 70 | 2:41 | 77 | Unknown | 0.064 | 0.0496 | 0 | 0.0057 | 0.41016 | 0.0016
2018-08-02 | 13 | c | 71 | 2:42 | 78 | Unknown | 0.064 | 0.0496 | 0 | 0.005 | 0.41016 | 0.0016
2018-08-16 | 1 | a | 72 | 2:43 | 79 | Unknown | 0.064 | 0.0496 | 0.004 | 0.0206 | 0.41016 | 0.017663
2018-08-16 | 1 | b | 73 | 2:44 | 80 | Unknown | 0.064 | 0.0495 | 0.003 | 0.0173 | 0.41016 | 0.013635
2018-08-16 | 1 | c | 74 | 2:45 | 81 | Unknown | 0.064 | 0.0495 | 0.003 | 0.0167 | 0.41016 | 0.013635
2018-08-16 | 2 | a | 75 | 2:46 | 82 | Unknown | 0.065 | 0.0498 | 0.003 | 0.0167 | 0.41666 | 0.013635
2018-08-16 | 2 | b | 76 | 2:47 | 83 | Unknown | 0.065 | 0.0499 | 0.003 | 0.016 | 0.41666 | 0.013635
2018-08-16 | 2 | c | 78 | 2:48 | 84 | Unknown | 0.065 | 0.0499 | 0.004 | 0.0189 | 0.41666 | 0.017663
2018-08-16 | 3 | a | 79 | 2:49 | 85 | Unknown | 0.065 | 0.05 | 0.004 | 0.0194 | 0.41666 | 0.017663
2018-08-16 | 3 | b | 80 | 2:50 | 86 | Unknown | 0.065 | 0.0499 | 0.003 | 0.0175 | 0.41666 | 0.013635
2018-08-16 | 3 | c | 81 | 2:51 | 87 | Unknown | 0.065 | 0.0499 | 0.003 | 0.0169 | 0.41666 | 0.013635
2018-08-16 | 4 | a | 82 | 2:52 | 88 | Unknown | 0.065 | 0.0499 | 0.002 | 0.0116 | 0.41666 | 0.0096155
2018-08-16 | 4 | b | 83 | 2:53 | 89 | Unknown | 0.064 | 0.0497 | 0.002 | 0.0121 | 0.41016 | 0.0096155
2018-08-16 | 4 | c | 84 | 2:54 | 90 | Unknown | 0.065 | 0.0499 | 0.003 | 0.0147 | 0.41666 | 0.013635
2018-08-16 | 5 | a | 85 | 2:55 | 91 | Unknown | 0.064 | 0.0494 | 0.002 | 0.0093 | 0.41016 | 0.0096155
2018-08-16 | 5 | b | 86 | 2:56 | 92 | Unknown | 0.064 | 0.0495 | 0.001 | 0.0079 | 0.41016 | 0.0056037
2018-08-16 | 5 | c | 87 | 2:57 | 93 | Unknown | 0.064 | 0.0496 | 0.001 | 0.0068 | 0.41016 | 0.0056037
2018-08-16 | 6 | a | 89 | 2:58 | 94 | Unknown | 0.064 | 0.0496 | 0 | 0.0044 | 0.41016 | 0.0016
2018-08-16 | 6 | b | 90 | 2:59 | 95 | Unknown | 0.064 | 0.0496 | 0.022 | 0.0931 | 0.41016 | 0.091558
2018-08-16 | 6 | c | 91 | 2:60 | 96 | Unknown | 0.064 | 0.0497 | 0.004 | 0.0171 | 0.41016 | 0.017663
2018-08-16 | 7 | a | 92 | 3:1 | 97 | Unknown | 0.064 | 0.0497 | 0 | 0.0036 | 0.41016 | 0.0016
2018-08-16 | 7 | b | 93 | 3:2 | 98 | Unknown | 0.064 | 0.0497 | 0 | 0.0018 | 0.41016 | 0.0016
2018-08-16 | 7 | c | 94 | 3:3 | 99 | Unknown | 0.064 | 0.0496 | 0 | 0.0033 | 0.41016 | 0.0016
2018-08-16 | 8 | a | 95 | 3:4 | 100 | Unknown | 0.064 | 0.0494 | 0 | 0.0022 | 0.41016 | 0.0016
2018-08-16 | 8 | b | 96 | 3:5 | 101 | Unknown | 0.064 | 0.0495 | 0.001 | 0.0075 | 0.41016 | 0.0056037
2018-08-16 | 8 | c | 97 | 3:6 | 102 | Unknown | 0.064 | 0.0495 | 0 | 0.0042 | 0.41016 | 0.0016
2018-08-16 | 9 | a | 98 | 3:7 | 103 | Unknown | 0.064 | 0.0495 | 0.002 | 0.0112 | 0.41016 | 0.0096155
2018-08-16 | 9 | b | 100 | 3:8 | 104 | Unknown | 0.065 | 0.0501 | 0.002 | 0.0093 | 0.41666 | 0.0096155
2018-08-16 | 9 | c | 101 | 3:9 | 105 | Unknown | 0.064 | 0.0496 | 0.001 | 0.0085 | 0.41016 | 0.0056037
2018-08-16 | 10 | a | 102 | 3:10 | 106 | Unknown | 0.064 | 0.0495 | 0.002 | 0.0126 | 0.41016 | 0.0096155
2018-08-16 | 10 | b | 103 | 3:11 | 107 | Unknown | 0.064 | 0.0494 | 0.002 | 0.0124 | 0.41016 | 0.0096155
2018-08-16 | 10 | c | 104 | 3:12 | 108 | Unknown | 0.064 | 0.0495 | 0.001 | 0.0087 | 0.41016 | 0.0056037
2018-08-16 | 11 | a | 105 | 3:13 | 109 | Unknown | 0.064 | 0.0498 | 0.001 | 0.0063 | 0.41016 | 0.0056037
2018-08-16 | 11 | b | 106 | 3:14 | 110 | Unknown | 0.064 | 0.0496 | 0.001 | 0.0084 | 0.41016 | 0.0056037
2018-08-16 | 11 | c | 107 | 3:15 | 111 | Unknown | 0.064 | 0.0494 | 0.001 | 0.0072 | 0.41016 | 0.0056037
2018-08-16 | 12 | a | 108 | 3:16 | 112 | Unknown | 0.065 | 0.0498 | 0.001 | 0.0047 | 0.41666 | 0.0056037
2018-08-16 | 12 | b | 109 | 3:17 | 113 | Unknown | 0.064 | 0.0496 | 0.001 | 0.005 | 0.41016 | 0.0056037
2018-08-16 | 12 | c | 111 | 3:18 | 114 | Unknown | 0.064 | 0.0496 | 0 | 0.0031 | 0.41016 | 0.0016
2018-08-16 | 13 | a | 112 | 3:19 | 115 | Unknown | 0.064 | 0.0497 | 0 | 0.0012 | 0.41016 | 0.0016
2018-08-16 | 13 | b | 113 | 3:20 | 116 | Unknown | 0.064 | 0.0496 | 0 | 0.0007 | 0.41016 | 0.0016
2018-08-16 | 13 | c | 114 | 3:21 | 117 | Unknown | 0.064 | 0.0496 | 0 | 0.0013 | 0.41016 | 0.0016
2018-08-31 | 1 | a | 115 | 3:22 | 118 | Unknown | 0.064 | 0.0496 | 0.002 | 0.0125 | 0.41016 | 0.0096155
2018-08-31 | 1 | b | 116 | 3:23 | 119 | Unknown | 0.064 | 0.0496 | 0.003 | 0.0145 | 0.41016 | 0.013635
2018-08-31 | 1 | c | 117 | 3:24 | 120 | Unknown | 0.064 | 0.0496 | 0.003 | 0.0132 | 0.41016 | 0.013635
2018-08-31 | 2 | a | 118 | 3:25 | 121 | Unknown | 0.065 | 0.0498 | 0.003 | 0.0133 | 0.41666 | 0.013635
2018-08-31 | 2 | b | 119 | 3:26 | 122 | Unknown | 0.065 | 0.0498 | 0.07 | 0.2928 | 0.41666 | 0.30148
2018-08-31 | 2 | c | 120 | 3:27 | 123 | Unknown | 0.064 | 0.0496 | 0.003 | 0.0131 | 0.41016 | 0.013635
2018-08-31 | 3 | a | 122 | 3:28 | 124 | Unknown | 0.099 | 0.0667 | 0.003 | 0.0126 | 0.6375 | 0.013635
2018-08-31 | 3 | b | 123 | 3:29 | 125 | Unknown | 0.064 | 0.0492 | 0.002 | 0.0105 | 0.41016 | 0.0096155
2018-08-31 | 3 | c | 124 | 3:30 | 126 | Unknown | 0.064 | 0.0496 | 0.049 | 0.205 | 0.41016 | 0.20734
2018-08-31 | 4 | a | 125 | 3:31 | 127 | Unknown | 0.064 | 0.0497 | 0.003 | 0.0099 | 0.41016 | 0.013635
2018-08-31 | 4 | b | 126 | 3:32 | 128 | Unknown | 0.065 | 0.05 | 0.069 | 0.2868 | 0.41666 | 0.29692
2018-08-31 | 4 | c | 127 | 3:33 | 129 | Unknown | 0.064 | 0.0496 | 0.004 | 0.0132 | 0.41016 | 0.017663
2018-08-31 | 5 | a | 128 | 3:34 | 130 | Unknown | 0.064 | 0.0497 | 0.003 | 0.0072 | 0.41016 | 0.013635
2018-08-31 | 5 | b | 129 | 3:35 | 131 | Unknown | 0.064 | 0.0496 | 0.163 | 0.6751 | 0.41016 | 0.76148
2018-08-31 | 5 | c | 130 | 3:36 | 132 | Unknown | 0.064 | 0.0497 | 0.001 | -0.0028 | 0.41016 | 0.0056037
2018-08-31 | 6 | a | 131 | 3:37 | 133 | Unknown | 0.064 | 0.0496 | 0.443 | 1.8361 | 0.41016 | 2.5708
2018-08-31 | 6 | b | 134 | 3:38 | 134 | Unknown | 0.065 | 0.0496 | 0.14 | 0.5783 | 0.41666 | 0.64118
2018-08-31 | 6 | c | 135 | 3:39 | 135 | Unknown | 0.064 | 0.0496 | 0.038 | 0.1494 | 0.41016 | 0.15945
2018-08-31 | 7 | a | 136 | 3:40 | 136 | Unknown | 0.064 | 0.0496 | 0.003 | 0.0072 | 0.41016 | 0.013635
2018-08-31 | 7 | b | 137 | 3:41 | 137 | Unknown | 0.065 | 0.0501 | 0.002 | 0.0074 | 0.41666 | 0.0096155
2018-08-31 | 7 | c | 138 | 3:42 | 138 | Unknown | 0.064 | 0.0496 | 0.001 | 0.003 | 0.41016 | 0.0056037
2018-08-31 | 8 | a | 139 | 3:43 | 139 | Unknown | 0.064 | 0.0496 | 0.035 | 0.1434 | 0.41016 | 0.14656
2018-08-31 | 8 | b | 140 | 3:44 | 140 | Unknown | 0.065 | 0.0497 | 0.028 | 0.1137 | 0.41666 | 0.11677
2018-08-31 | 8 | c | 141 | 3:45 | 141 | Unknown | 0.064 | 0.0495 | 0.001 | 0.0065 | 0.41016 | 0.0056037
2018-08-31 | 9 | a | 142 | 3:46 | 142 | Unknown | 0.064 | 0.0496 | 0.023 | 0.0969 | 0.41016 | 0.09574
2018-08-31 | 9 | b | 144 | 3:47 | 143 | Unknown | 0.064 | 0.0493 | 0.12 | 0.5024 | 0.41016 | 0.54006
2018-08-31 | 9 | c | 145 | 3:48 | 144 | Unknown | 0.064 | 0.0494 | 0.145 | 0.6007 | 0.41016 | 0.66696
2018-08-31 | 10 | a | 146 | 3:49 | 145 | Unknown | 0.064 | 0.0495 | 0.157 | 0.6508 | 0.41016 | 0.72968
2018-08-31 | 10 | b | 147 | 3:50 | 146 | Unknown | 0.064 | 0.0494 | 0.036 | 0.1466 | 0.41016 | 0.15085
2018-08-31 | 10 | c | 148 | 3:51 | 147 | Unknown | 0.064 | 0.0495 | 0.149 | 0.6197 | 0.41016 | 0.68774
2018-08-31 | 11 | a | 149 | 3:52 | 148 | Unknown | 0.064 | 0.0494 | 0.304 | 1.2595 | 0.41016 | 1.593
2018-08-31 | 11 | b | 150 | 3:53 | 149 | Unknown | 0.064 | 0.0494 | 0.101 | 0.4113 | 0.41016 | 0.44701
2018-08-31 | 11 | c | 151 | 3:54 | 150 | Unknown | 0.064 | 0.0494 | 0.102 | 0.4199 | 0.41016 | 0.45183
2018-08-31 | 12 | a | 152 | 3:55 | 151 | Unknown | 0.064 | 0.0496 | 0.195 | 0.8072 | 0.41016 | 0.93601
2018-08-31 | 12 | b | 153 | 3:56 | 152 | Unknown | 0.064 | 0.0496 | 0.002 | 0.0005 | 0.41016 | 0.0096155
2018-08-31 | 12 | c | 155 | 3:57 | 153 | Unknown | 0.064 | 0.0496 | 0 | -0.0019 | 0.41016 | 0.0016
2018-08-31 | 13 | a | 156 | 3:58 | 154 | Unknown | 0.064 | 0.0496 | -0.001 | -0.0059 | 0.41016 | -0.0023955
2018-08-31 | 13 | b | 157 | 3:59 | 155 | Unknown | 0.064 | 0.0496 | 0 | -0.0032 | 0.41016 | 0.0016
2018-08-31 | 13 | c | 158 | 3:60 | 156 | Unknown | 0.076 | 0.0551 | ??? | ??? | 0.48811 | #VALUE!

## Final Results
SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | NH3 mg/L | P-PO4 mg/L
2018-07-18 | 13 | a |  | 0.0056037
2018-07-18 | 13 | b |  | 0.0016
2018-07-18 | 13 | c |  | 0.0016
2018-08-02 | 1 | a |  | 0.017663
2018-08-02 | 1 | b |  | 0.013635
2018-08-02 | 1 | c |  | 0.0096155
2018-08-02 | 2 | a |  | 0.017663
2018-08-02 | 2 | b |  | 0.013635
2018-08-02 | 2 | c |  | 0.017663
2018-08-02 | 3 | a |  | 0.017663
2018-08-02 | 3 | b |  | 0.017663
2018-08-02 | 3 | c |  | 0.0217
2018-08-02 | 4 | a |  | 0.0217
2018-08-02 | 4 | b |  | 0.0096155
2018-08-02 | 4 | c |  | 0.013635
2018-08-02 | 5 | a |  | 0.0096155
2018-08-02 | 5 | b |  | 0.0096155
2018-08-02 | 5 | c |  | 0.0056037
2018-08-02 | 6 | a |  | 0.0056037
2018-08-02 | 6 | b |  | 0.0056037
2018-08-02 | 6 | c |  | 0.0056037
2018-08-02 | 7 | a |  | 0.0056037
2018-08-02 | 7 | b |  | 0.0056037
2018-08-02 | 7 | c |  | 0.0056037
2018-08-02 | 8 | a |  | 0.0056037
2018-08-02 | 8 | b |  | 0.0056037
2018-08-02 | 8 | c |  | 0.0056037
2018-08-02 | 9 | a |  | 0.0056037
2018-08-02 | 9 | b |  | 0.0096155
2018-08-02 | 9 | c |  | 0.0096155
2018-08-02 | 10 | a |  | 0.0096155
2018-08-02 | 10 | b |  | 0.0096155
2018-08-02 | 10 | c |  | 0.0096155
2018-08-02 | 11 | a |  | 0.013635
2018-08-02 | 11 | b |  | 0.0096155
2018-08-02 | 11 | c |  | 0.013635
2018-08-02 | 12 | a |  | 0.0096155
2018-08-02 | 12 | b |  | 0.0096155
2018-08-02 | 12 | c |  | 0.0056037
2018-08-02 | 13 | a |  | 0.0016
2018-08-02 | 13 | b |  | 0.0016
2018-08-02 | 13 | c |  | 0.0016
2018-08-16 | 1 | a |  | 0.017663
2018-08-16 | 1 | b |  | 0.013635
2018-08-16 | 1 | c |  | 0.013635
2018-08-16 | 2 | a |  | 0.013635
2018-08-16 | 2 | b |  | 0.013635
2018-08-16 | 2 | c |  | 0.017663
2018-08-16 | 3 | a |  | 0.017663
2018-08-16 | 3 | b |  | 0.013635
2018-08-16 | 3 | c |  | 0.013635
2018-08-16 | 4 | a |  | 0.0096155
2018-08-16 | 4 | b |  | 0.0096155
2018-08-16 | 4 | c |  | 0.013635
2018-08-16 | 5 | a |  | 0.0096155
2018-08-16 | 5 | b |  | 0.0056037
2018-08-16 | 5 | c |  | 0.0056037
2018-08-16 | 6 | a |  | 0.0016
2018-08-16 | 6 | b |  | 0.091558
2018-08-16 | 6 | c |  | 0.017663
2018-08-16 | 7 | a |  | 0.0016
2018-08-16 | 7 | b |  | 0.0016
2018-08-16 | 7 | c |  | 0.0016
2018-08-16 | 8 | a |  | 0.0016
2018-08-16 | 8 | b |  | 0.0056037
2018-08-16 | 8 | c |  | 0.0016
2018-08-16 | 9 | a |  | 0.0096155
2018-08-16 | 9 | b |  | 0.0096155
2018-08-16 | 9 | c |  | 0.0056037
2018-08-16 | 10 | a |  | 0.0096155
2018-08-16 | 10 | b |  | 0.0096155
2018-08-16 | 10 | c |  | 0.0056037
2018-08-16 | 11 | a |  | 0.0056037
2018-08-16 | 11 | b |  | 0.0056037
2018-08-16 | 11 | c |  | 0.0056037
2018-08-16 | 12 | a |  | 0.0056037
2018-08-16 | 12 | b |  | 0.0056037
2018-08-16 | 12 | c |  | 0.0016
2018-08-16 | 13 | a |  | 0.0016
2018-08-16 | 13 | b |  | 0.0016
2018-08-16 | 13 | c |  | 0.0016
2018-08-31 | 1 | a |  | 0.0096155
2018-08-31 | 1 | b |  | 0.013635
2018-08-31 | 1 | c |  | 0.013635
2018-08-31 | 2 | a |  | 0.013635
2018-08-31 | 2 | b |  | 0.30148
2018-08-31 | 2 | c |  | 0.013635
2018-08-31 | 3 | a |  | 0.013635
2018-08-31 | 3 | b |  | 0.0096155
2018-08-31 | 3 | c |  | 0.20734
2018-08-31 | 4 | a |  | 0.013635
2018-08-31 | 4 | b |  | 0.29692
2018-08-31 | 4 | c |  | 0.017663
2018-08-31 | 5 | a |  | 0.013635
2018-08-31 | 5 | b |  | 0.76148
2018-08-31 | 5 | c |  | 0.0056037
2018-08-31 | 6 | a |  | 2.5708
2018-08-31 | 6 | b |  | 0.64118
2018-08-31 | 6 | c |  | 0.15945
2018-08-31 | 7 | a |  | 0.013635
2018-08-31 | 7 | b |  | 0.0096155
2018-08-31 | 7 | c |  | 0.0056037
2018-08-31 | 8 | a |  | 0.14656
2018-08-31 | 8 | b |  | 0.11677
2018-08-31 | 8 | c |  | 0.0056037
2018-08-31 | 9 | a |  | 0.09574
2018-08-31 | 9 | b |  | 0.54006
2018-08-31 | 9 | c |  | 0.66696
2018-08-31 | 10 | a |  | 0.72968
2018-08-31 | 10 | b |  | 0.15085
2018-08-31 | 10 | c |  | 0.68774
2018-08-31 | 11 | a |  | 1.593
2018-08-31 | 11 | b |  | 0.44701
2018-08-31 | 11 | c |  | 0.45183
2018-08-31 | 12 | a |  | 0.93601
2018-08-31 | 12 | b |  | 0.0096155
2018-08-31 | 12 | c |  | 0.0016
2018-08-31 | 13 | a |  | -0.0023955
2018-08-31 | 13 | b |  | 0.0016
2018-08-31 | 13 | c |  | #VALUE!