## Results and Sample ID
[5 empty] | Run Name: | LTREB_NO3-AP2_2019~03~07
Run Date:
[5 empty] | Operator: | RE
|  |  |  | Sample Info |  |  |  | FASPac NO3-NO2 |  |  | Calibrant | Conc., mg/L
Row | SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | Peak# | Position | Identifier | Type | Raw Ht | Corr Ht | mg/L |  | C1 | 0.001
|  |  |  | 1 | 1:1 | SYNC | SYNC | 0.382 | 0.38 | 0.5 |  | C2 | 0.005
|  |  |  | 2 | 1:2 | CO | Carry over | 0.003 | 0.001 | -0.002 |  | C3 | 0.01
|  |  |  | 3 | 1:3 | W | Wash | 0.002 | 0 | -0.003 |  | C4 | 0.05
|  |  |  | 4 | 1:8 | NO3 | Unknown | 0.082 | 0.08 | 0.1 |  | C5 | 0.1
|  |  |  | 5 | 1:12 | NOX% | NO3 check | 0.08 | 0.078 | 0.097 |  | C6 | 0.5
[6 empty] | NO3 Efficiency |  |  |  | 103.56%
|  |  |  | 6 | 1:4 | C1 | Calibrant | 0.003 | 0.001 | -0.001
|  |  |  | 7 | 1:4 | C1 | Calibrant | 0.003 | 0.001 | -0.002
|  |  |  | 8 | 1:4 | C1 | Calibrant | 0.003 | 0.001 | -0.002
|  |  |  | 9 | 1:5 | C2 | Calibrant | 0.009 | 0.007 | 0.006
|  |  |  | 10 | 1:5 | C2 | Calibrant | 0.009 | 0.007 | 0.006
|  |  |  | 11 | 1:5 | C2 | Calibrant | 0.009 | 0.007 | 0.006
|  |  |  | 12 | 1:6 | C3 | Calibrant | 0.014 | 0.012 | 0.012
|  |  |  | 13 | 1:6 | C3 | Calibrant | 0.014 | 0.012 | 0.012
|  |  |  | 14 | 1:6 | C3 | Calibrant | 0.014 | 0.012 | 0.012
|  |  |  | 15 | 1:7 | C4 | Calibrant | 0.045 | 0.042 | 0.051
|  |  |  | 16 | 1:7 | C4 | Calibrant | 0.045 | 0.042 | 0.051
|  |  |  | 17 | 1:7 | C4 | Calibrant | 0.045 | 0.042 | 0.051
|  |  |  | 18 | 1:8 | C5 | Calibrant | 0.083 | 0.08 | 0.1
|  |  |  | 19 | 1:8 | C5 | Calibrant | 0.082 | 0.079 | 0.099
|  |  |  | 20 | 1:8 | C5 | Calibrant | 0.082 | 0.079 | 0.099
|  |  |  | 21 | 1:9 | C6 | Calibrant | 0.382 | 0.378 | 0.497
|  |  |  | 22 | 1:9 | C6 | Calibrant | 0.386 | 0.382 | 0.502
|  |  |  | 23 | 1:9 | C6 | Calibrant | 0.386 | 0.382 | 0.501
|  |  |  | 24 | 1:10 | CC1 | QC 2 | 0.057 | 0.052 | 0.064
|  |  |  | 25 | 1:11 | CC2 | Unknown | 0.315 | 0.312 | 0.407
|  |  |  | 26 | 1:3 | W | Wash | 0.003 | 0 | -0.003
|  |  |  | 27 | 1:8 | NO3 | Unknown | 0.082 | 0.079 | 0.098
|  |  |  | 28 | 1:12 | NOX% | NO3 check | 0.081 | 0.078 | 0.097
[6 empty] | NO3 Efficiency |  |  |  | 101.67%
| 2018-09-14 | 1 | a | 29 | 1:25 | 1 | Unknown | 0.043 | 0.04 | 0.048
| 2018-09-14 | 1 | b | 30 | 1:26 | 2 | Unknown | 0.044 | 0.041 | 0.049
| 2018-09-14 | 1 | c | 31 | 1:27 | 3 | Unknown | 0.042 | 0.039 | 0.047
| 2018-09-14 | 2 | a | 32 | 1:28 | 4 | Unknown | 0.031 | 0.029 | 0.034
| 2018-09-14 | 2 | b | 33 | 1:29 | 5 | Unknown | 0.031 | 0.029 | 0.034
| 2018-09-14 | 2 | c | 34 | 1:30 | 6 | Unknown | 0.032 | 0.029 | 0.035
| 2018-09-14 | 3 | a | 36 | 1:31 | 7 | Unknown | 0.062 | 0.06 | 0.074
| 2018-09-14 | 3 | b | 37 | 1:32 | 8 | Unknown | 0.067 | 0.065 | 0.081
| 2018-09-14 | 3 | c | 38 | 1:33 | 9 | Unknown | 0.065 | 0.063 | 0.078
| 2018-09-14 | 4 | a | 39 | 1:34 | 10 | Unknown | 0.028 | 0.025 | 0.029
| 2018-09-14 | 4 | b | 40 | 1:35 | 11 | Unknown | 0.029 | 0.026 | 0.031
| 2018-09-14 | 4 | c | 41 | 1:36 | 12 | Unknown | 0.027 | 0.025 | 0.029
| 2018-09-14 | 5 | a | 42 | 1:37 | 13 | Unknown | 0.054 | 0.052 | 0.063
| 2018-09-14 | 5 | b | 43 | 1:38 | 14 | Unknown | 0.042 | 0.04 | 0.048
| 2018-09-14 | 5 | c | 44 | 1:39 | 15 | Unknown | 0.05 | 0.047 | 0.057
| 2018-09-14 | 6 | a | 45 | 1:40 | 16 | Unknown | 0.106 | 0.103 | 0.13
| 2018-09-14 | 6 | b | 47 | 1:41 | 17 | Unknown | 0.118 | 0.115 | 0.146
| 2018-09-14 | 6 | c | 48 | 1:42 | 18 | Unknown | 0.114 | 0.111 | 0.14
| 2018-09-14 | 7 | a | 49 | 1:43 | 19 | Unknown | 0.108 | 0.105 | 0.132
| 2018-09-14 | 7 | b | 50 | 1:44 | 20 | Unknown | 0.117 | 0.114 | 0.145
| 2018-09-14 | 7 | c | 51 | 1:45 | 21 | Unknown | 0.117 | 0.114 | 0.144
| 2018-09-14 | 8 | a | 52 | 1:46 | 22 | Unknown | 0.11 | 0.107 | 0.135
| 2018-09-14 | 8 | b | 53 | 1:47 | 23 | Unknown | 0.093 | 0.091 | 0.114
| 2018-09-14 | 8 | c | 54 | 1:48 | 24 | Unknown | 0.092 | 0.089 | 0.112
| 2018-09-14 | 9 | a | 55 | 1:49 | 25 | Unknown | 0.009 | 0.006 | 0.005
| 2018-09-14 | 9 | b | 56 | 1:50 | 26 | Unknown | 0.008 | 0.005 | 0.004
| 2018-09-14 | 9 | c | 58 | 1:51 | 27 | Unknown | 0.007 | 0.005 | 0.003
| 2018-09-14 | 10 | a | 59 | 1:52 | 28 | Unknown | 0.005 | 0.003 | 0.001
| 2018-09-14 | 10 | b | 60 | 1:53 | 29 | Unknown | 0.005 | 0.003 | 0
| 2018-09-14 | 10 | c | 63 | 1:54 | 30 | Unknown | 0.007 | 0.004 | 0.002
| 2018-09-14 | 11 | a | 64 | 1:55 | 31 | Unknown | 0.006 | 0.003 | 0.001
| 2018-09-14 | 11 | b | 65 | 1:56 | 32 | Unknown | 0.007 | 0.004 | 0.002
| 2018-09-14 | 11 | c | 66 | 1:57 | 33 | Unknown | 0.006 | 0.003 | 0
| 2018-09-14 | 12 | a | 67 | 1:58 | 34 | Unknown | 0.011 | 0.008 | 0.007
| 2018-09-14 | 12 | b | 69 | 1:59 | 35 | Unknown | 0.011 | 0.008 | 0.007
| 2018-09-14 | 12 | c | 72 | 1:60 | 36 | Unknown | 0.011 | 0.007 | 0.006
| 2018-09-14 | 13 | a | 73 | 2:1 | 37 | Unknown | 0.005 | 0.002 | -0.001
| 2018-09-14 | 13 | b | 74 | 2:2 | 38 | Unknown | 0.005 | 0.002 | -0.001
| 2018-09-14 | 13 | c | 75 | 2:3 | 39 | Unknown | 0.005 | 0.002 | -0.001
| 2018-09-27 | 1 | a | 76 | 2:4 | 40 | Unknown | 0.053 | 0.049 | 0.06
| 2018-09-27 | 1 | b | 77 | 2:5 | 41 | Unknown | 0.059 | 0.055 | 0.067
| 2018-09-27 | 1 | c | 78 | 2:6 | 42 | Unknown | 0.055 | 0.051 | 0.063
| 2018-09-27 | 2 | a | 80 | 2:7 | 43 | Unknown | 0.033 | 0.029 | 0.034
| 2018-09-27 | 2 | b | 81 | 2:8 | 44 | Unknown | 0.034 | 0.03 | 0.036
| 2018-09-27 | 2 | c | 82 | 2:9 | 45 | Unknown | 0.032 | 0.028 | 0.033
| 2018-09-27 | 3 | a | 83 | 2:10 | 46 | Unknown | 0.089 | 0.085 | 0.106
| 2018-09-27 | 3 | b | 84 | 2:11 | 47 | Unknown | 0.081 | 0.076 | 0.095
| 2018-09-27 | 3 | c | 85 | 2:12 | 48 | Unknown | 0.078 | 0.074 | 0.092
| 2018-09-27 | 4 | a | 86 | 2:13 | 49 | Unknown | 0.038 | 0.033 | 0.039
| 2018-09-27 | 4 | b | 87 | 2:14 | 50 | Unknown | 0.038 | 0.033 | 0.04
| 2018-09-27 | 4 | c | 91 | 2:15 | 51 | Unknown | 0.038 | 0.033 | 0.039
| 2018-09-27 | 5 | a | 92 | 2:16 | 52 | Unknown | 0.05 | 0.045 | 0.055
| 2018-09-27 | 5 | b | 93 | 2:17 | 53 | Unknown | 0.051 | 0.046 | 0.056
| 2018-09-27 | 5 | c | 94 | 2:18 | 54 | Unknown | 0.057 | 0.052 | 0.064
| 2018-09-27 | 6 | a | 95 | 2:19 | 55 | Unknown | 0.128 | 0.123 | 0.156
| 2018-09-27 | 6 | b | 96 | 2:20 | 56 | Unknown | 0.132 | 0.127 | 0.161
| 2018-09-27 | 6 | c | 97 | 2:21 | 57 | Unknown | 0.139 | 0.134 | 0.17
| 2018-09-27 | 7 | a | 98 | 2:22 | 58 | Unknown | 0.121 | 0.116 | 0.147
| 2018-09-27 | 7 | b | 99 | 2:23 | 59 | Unknown | 0.128 | 0.123 | 0.156
| 2018-09-27 | 7 | c | 100 | 2:24 | 60 | Unknown | 0.11 | 0.106 | 0.133
| 2018-09-27 | 8 | a | 102 | 2:25 | 61 | Unknown | 0.091 | 0.086 | 0.108
| 2018-09-27 | 8 | b | 103 | 2:26 | 62 | Unknown | 0.089 | 0.084 | 0.106
| 2018-09-27 | 8 | c | 104 | 2:27 | 63 | Unknown | 0.094 | 0.09 | 0.112
| 2018-09-27 | 9 | a | 105 | 2:28 | 64 | Unknown | 0.008 | 0.003 | 0.001
| 2018-09-27 | 9 | b | 106 | 2:29 | 65 | Unknown | 0.007 | 0.003 | 0.001
| 2018-09-27 | 9 | c | 107 | 2:30 | 66 | Unknown | 0.008 | 0.004 | 0.001
| 2018-09-27 | 10 | a | 108 | 2:31 | 67 | Unknown | 0.009 | 0.005 | 0.004
| 2018-09-27 | 10 | b | 109 | 2:32 | 68 | Unknown | 0.009 | 0.005 | 0.004
| 2018-09-27 | 10 | c | 110 | 2:33 | 69 | Unknown | 0.008 | 0.005 | 0.003
| 2018-09-27 | 11 | a | 111 | 2:34 | 70 | Unknown | 0.008 | 0.005 | 0.003
| 2018-09-27 | 11 | b | 113 | 2:35 | 71 | Unknown | 0.007 | 0.004 | 0.002
| 2018-09-27 | 11 | c | 114 | 2:36 | 72 | Unknown | 0.007 | 0.004 | 0.002
| 2018-09-27 | 12 | a | 115 | 2:37 | 73 | Unknown | 0.013 | 0.009 | 0.009
| 2018-09-27 | 12 | b | 116 | 2:38 | 74 | Unknown | 0.013 | 0.009 | 0.008
| 2018-09-27 | 12 | c | 117 | 2:39 | 75 | Unknown | 0.014 | 0.01 | 0.01
| 2018-09-27 | 13 | a | 118 | 2:40 | 76 | Unknown | 0.007 | 0.003 | 0.001
| 2018-09-27 | 13 | b | 119 | 2:41 | 77 | Unknown | 0.006 | 0.002 | 0
| 2018-09-27 | 13 | c | 120 | 2:42 | 78 | Unknown | 0.006 | 0.002 | 0
| 2018-10-11 | 1 | a | 121 | 2:43 | 79 | Unknown | 0.032 | 0.028 | 0.033
| 2018-10-11 | 1 | b | 122 | 2:44 | 80 | Unknown | 0.032 | 0.028 | 0.033
| 2018-10-11 | 1 | c | 124 | 2:45 | 81 | Unknown | 0.032 | 0.028 | 0.033
| 2018-10-11 | 2 | a | 127 | 2:46 | 82 | Unknown | 0.027 | 0.023 | 0.026
| 2018-10-11 | 2 | b | 128 | 2:47 | 83 | Unknown | 0.026 | 0.022 | 0.025
| 2018-10-11 | 2 | c | 129 | 2:48 | 84 | Unknown | 0.026 | 0.022 | 0.025
| 2018-10-11 | 3 | a | 130 | 2:49 | 85 | Unknown | 0.061 | 0.057 | 0.07
| 2018-10-11 | 3 | b | 131 | 2:50 | 86 | Unknown | 0.057 | 0.053 | 0.065
| 2018-10-11 | 3 | c | 132 | 2:51 | 87 | Unknown | 0.066 | 0.062 | 0.076
| 2018-10-11 | 4 | a | 133 | 2:52 | 88 | Unknown | 0.046 | 0.042 | 0.05
| 2018-10-11 | 4 | b | 135 | 2:53 | 89 | Unknown | 0.046 | 0.042 | 0.051
| 2018-10-11 | 4 | c | 136 | 2:54 | 90 | Unknown | 0.051 | 0.047 | 0.058
| 2018-10-11 | 5 | a | 137 | 2:55 | 91 | Unknown | 0.063 | 0.059 | 0.072
| 2018-10-11 | 5 | b | 138 | 2:56 | 92 | Unknown | 0.065 | 0.061 | 0.075
| 2018-10-11 | 5 | c | 139 | 2:57 | 93 | Unknown | 0.063 | 0.059 | 0.073
| 2018-10-11 | 6 | a | 140 | 2:58 | 94 | Unknown | 0.102 | 0.098 | 0.123
| 2018-10-11 | 6 | b | 141 | 2:59 | 95 | Unknown | 0.094 | 0.09 | 0.113
| 2018-10-11 | 6 | c | 144 | 2:60 | 96 | Unknown | 0.088 | 0.083 | 0.104
| 2018-10-11 | 7 | a | 146 | 3:1 | 97 | Unknown | 0.118 | 0.114 | 0.144
| 2018-10-11 | 7 | b | 147 | 3:2 | 98 | Unknown | 0.1 | 0.096 | 0.12
| 2018-10-11 | 7 | c | 148 | 3:3 | 99 | Unknown | 0.114 | 0.11 | 0.138
| 2018-10-11 | 8 | a | 149 | 3:4 | 100 | Unknown | 0.11 | 0.105 | 0.133
| 2018-10-11 | 8 | b | 150 | 3:5 | 101 | Unknown | 0.099 | 0.095 | 0.119
| 2018-10-11 | 8 | c | 151 | 3:6 | 102 | Unknown | 0.095 | 0.091 | 0.114
| 2018-10-11 | 9 | a | 152 | 3:7 | 103 | Unknown | 0.047 | 0.043 | 0.052
| 2018-10-11 | 9 | b | 153 | 3:8 | 104 | Unknown | 0.049 | 0.045 | 0.054
| 2018-10-11 | 9 | c | 154 | 3:9 | 105 | Unknown | 0.049 | 0.045 | 0.055
| 2018-10-11 | 10 | a | 155 | 3:10 | 106 | Unknown | 0.021 | 0.017 | 0.019
| 2018-10-11 | 10 | b | 157 | 3:11 | 107 | Unknown | 0.02 | 0.016 | 0.018
| 2018-10-11 | 10 | c | 158 | 3:12 | 108 | Unknown | 0.017 | 0.014 | 0.014
| 2018-10-11 | 11 | a | 159 | 3:13 | 109 | Unknown | 0.018 | 0.014 | 0.015
| 2018-10-11 | 11 | b | 160 | 3:14 | 110 | Unknown | 0.013 | 0.009 | 0.009
| 2018-10-11 | 11 | c | 161 | 3:15 | 111 | Unknown | 0.017 | 0.013 | 0.014
| 2018-10-11 | 12 | a | 162 | 3:16 | 112 | Unknown | 0.019 | 0.015 | 0.017
| 2018-10-11 | 12 | b | 163 | 3:17 | 113 | Unknown | 0.018 | 0.014 | 0.015
| 2018-10-11 | 12 | c | 164 | 3:18 | 114 | Unknown | 0.019 | 0.016 | 0.017
| 2018-10-11 | 13 | a | 165 | 3:19 | 115 | Unknown | 0.009 | 0.005 | 0.004
| 2018-10-11 | 13 | b | 166 | 3:20 | 116 | Unknown | 0.009 | 0.005 | 0.003
| 2018-10-11 | 13 | c | 168 | 3:21 | 117 | Unknown | 0.008 | 0.005 | 0.003
| 2018-11-02 | 1 | a | 169 | 3:22 | 118 | Unknown | 0.027 | 0.023 | 0.026
| 2018-11-02 | 1 | b | 170 | 3:23 | 119 | Unknown | 0.025 | 0.021 | 0.023
| 2018-11-02 | 1 | c | 171 | 3:24 | 120 | Unknown | 0.026 | 0.023 | 0.026
| 2018-11-02 | 2 | a | 172 | 3:25 | 121 | Unknown | 0.013 | 0.009 | 0.008
| 2018-11-02 | 2 | b | 173 | 3:26 | 122 | Unknown | 0.011 | 0.007 | 0.006
| 2018-11-02 | 2 | c | 174 | 3:27 | 123 | Unknown | 0.012 | 0.008 | 0.007
| 2018-11-02 | 3 | a | 175 | 3:28 | 124 | Unknown | 0.051 | 0.047 | 0.057
| 2018-11-02 | 3 | b | 176 | 3:29 | 125 | Unknown | 0.048 | 0.044 | 0.054
| 2018-11-02 | 3 | c | 177 | 3:30 | 126 | Unknown | 0.051 | 0.047 | 0.057
| 2018-11-02 | 4 | a | 179 | 3:31 | 127 | Unknown | 0.045 | 0.041 | 0.05
| 2018-11-02 | 4 | b | 180 | 3:32 | 128 | Unknown | 0.045 | 0.04 | 0.049
| 2018-11-02 | 4 | c | 181 | 3:33 | 129 | Unknown | 0.047 | 0.043 | 0.052
| 2018-11-02 | 5 | a | 182 | 3:34 | 130 | Unknown | 0.06 | 0.056 | 0.069
| 2018-11-02 | 5 | b | 183 | 3:35 | 131 | Unknown | 0.056 | 0.052 | 0.063
| 2018-11-02 | 5 | c | 184 | 3:36 | 132 | Unknown | 0.057 | 0.053 | 0.065
| 2018-11-02 | 6 | a | 187 | 3:37 | 133 | Unknown | 0.108 | 0.103 | 0.13
| 2018-11-02 | 6 | b | 188 | 3:38 | 134 | Unknown | 0.102 | 0.098 | 0.123
| 2018-11-02 | 6 | c | 190 | 3:39 | 135 | Unknown | 0.104 | 0.099 | 0.125
| 2018-11-02 | 7 | a | 191 | 3:40 | 136 | Unknown | 0.098 | 0.093 | 0.117
| 2018-11-02 | 7 | b | 192 | 3:41 | 137 | Unknown | 0.093 | 0.089 | 0.112
| 2018-11-02 | 7 | c | 193 | 3:42 | 138 | Unknown | 0.105 | 0.1 | 0.127
| 2018-11-02 | 8 | a | 194 | 3:43 | 139 | Unknown | 0.097 | 0.093 | 0.117
| 2018-11-02 | 8 | b | 195 | 3:44 | 140 | Unknown | 0.1 | 0.096 | 0.12
| 2018-11-02 | 8 | c | 196 | 3:45 | 141 | Unknown | 0.093 | 0.089 | 0.111
| 2018-11-02 | 9 | a | 197 | 3:46 | 142 | Unknown | 0.044 | 0.04 | 0.048
| 2018-11-02 | 9 | b | 198 | 3:47 | 143 | Unknown | 0.042 | 0.038 | 0.045
| 2018-11-02 | 9 | c | 199 | 3:48 | 144 | Unknown | 0.046 | 0.042 | 0.051
| 2018-11-02 | 10 | a | 201 | 3:49 | 145 | Unknown | 0.051 | 0.047 | 0.057
| 2018-11-02 | 10 | b | 202 | 3:50 | 146 | Unknown | 0.05 | 0.047 | 0.057
| 2018-11-02 | 10 | c | 203 | 3:51 | 147 | Unknown | 0.052 | 0.049 | 0.059
| 2018-11-02 | 11 | a | 204 | 3:52 | 148 | Unknown | 0.05 | 0.046 | 0.056
| 2018-11-02 | 11 | b | 205 | 3:53 | 149 | Unknown | 0.05 | 0.046 | 0.056
| 2018-11-02 | 11 | c | 206 | 3:54 | 150 | Unknown | 0.05 | 0.046 | 0.056
| 2018-11-02 | 12 | a | 207 | 3:55 | 151 | Unknown | 0.044 | 0.04 | 0.049
| 2018-11-02 | 12 | b | 208 | 3:56 | 152 | Unknown | 0.043 | 0.04 | 0.048
| 2018-11-02 | 12 | c | 209 | 3:57 | 153 | Unknown | 0.043 | 0.039 | 0.048
| 2018-11-02 | 13 | a | 210 | 3:58 | 154 | Unknown | 0.028 | 0.024 | 0.028
| 2018-11-02 | 13 | b | 212 | 3:59 | 155 | Unknown | 0.028 | 0.024 | 0.028
| 2018-11-02 | 13 | c | 213 | 3:60 | 156 | Unknown | 0.028 | 0.024 | 0.028

## Std Curve Construction
1. Copy in Calibrant Information
[10 empty] | 2. Run Linear Regression | [8 empty] | 4. If quadratic exceeds 10%, run split-curve procedure.
Use X=Peak height; Y=Conc)
| Calibrant |  | Linear Fit |  | Quadratic |  | Split Curve
Calibrant | Conc. | Calibrant | Std Conc |  | Std Conc |  | Std Conc | [11 empty] | Plot low range calibrant data: |  |  | Enter coefficients below graph.
Id | (mg/L N) | peak hght (x) | (y-hat) | % error | (y-hat) | % error | (y-hat) | % error |  |  | All data - Linear Fit
C1 | 0.001 | 0.001 | -0.0028013 | -380.1291 | -0.0019181 | -291.8095 | 0.00085524 | -14.4763 |  | SUMMARY OUTPUT
C1 | 0.001 | 0.001 | -0.0028013 | -380.1291 | -0.0019181 | -291.8095 | 0.00085524 | -14.4763
C1 | 0.001 | 0.001 | -0.0028013 | -380.1291 | -0.0019181 | -291.8095 | 0.00085524 | -14.4763 |  | Regression Statistics
C2 | 0.005 | 0.007 | 0.0051383 | 2.7667 | 0.0057778 | 15.5553 | 0.00533 | 6.6003 |  | Multiple R | 0.99993
C2 | 0.005 | 0.007 | 0.0051383 | 2.7667 | 0.0057778 | 15.5553 | 0.00533 | 6.6003 |  | R Square | 0.99986
C2 | 0.005 | 0.007 | 0.0051383 | 2.7667 | 0.0057778 | 15.5553 | 0.00533 | 6.6003 |  | Adjusted R Square | 0.99985
C3 | 0.01 | 0.012 | 0.011755 | 17.5469 | 0.012197 | 21.9678 | 0.0097705 | -2.2947 |  | Standard Error | 0.0022568
C3 | 0.01 | 0.012 | 0.011755 | 17.5469 | 0.012197 | 21.9678 | 0.0097705 | -2.2947 |  | Observations | 18
C3 | 0.01 | 0.012 | 0.011755 | 17.5469 | 0.012197 | 21.9678 | 0.0097705 | -2.2947
C4 | 0.05 | 0.042 | 0.051453 | 2.9056 | 0.050822 | 1.6431 | 0.049997 | -0.005064 |  | ANOVA
C4 | 0.05 | 0.042 | 0.051453 | 2.9056 | 0.050822 | 1.6431 | 0.049997 | -0.005064 |  |  | df | SS | MS | F | Significance F
C4 | 0.05 | 0.042 | 0.051453 | 2.9056 | 0.050822 | 1.6431 | 0.049997 | -0.005064 |  | Regression | 1 | 0.56602 | 0.56602 | 111130.7986 | 3.6217e-32
C5 | 0.1 | 0.08 | 0.10174 | 1.7371 | 0.10002 | 0.01856 | 0.10103 | 1.0301 | Split | Residual | 16 | 8.1492e-05 | 5.0933e-06
C5 | 0.1 | 0.079 | 0.10041 | 0.41383 | 0.09872 | -1.28 | 0.099687 | -0.31343 |  | Total | 17 | 0.5661
C5 | 0.1 | 0.079 | 0.10041 | 0.41383 | 0.09872 | -1.28 | 0.099687 | -0.31343
C6 | 0.5 | 0.378 | 0.49607 | -0.78564 | 0.49638 | -0.72393 | 0.49647 | -0.70642 |  |  | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
C6 | 0.5 | 0.382 | 0.50136 | 0.27298 | 0.50183 | 0.3656 | 0.50171 | 0.3418 |  | Intercept | -0.0041246 | 0.00063421 | -6.5035 | 7.2766e-06 | -0.005469 | -0.0027801 | -0.005469 | -0.0027801 | Low curve coefficients:
C6 | 0.5 | 0.382 | 0.50136 | 0.27298 | 0.50183 | 0.3656 | 0.50171 | 0.3418 |  | X Variable 1 | 1.3233 | 0.0039695 | 333.3629 | 3.6217e-32 | 1.3149 | 1.3317 | 1.3149 | 1.3317 |  | intcpt | X coef. | X^2 coef.
[20 empty] | 0.0002 | 0.6423 | 12.937
[10 empty] | (Use X=Peak Ht.; Y=Conc) | [8 empty] | Plot High Range calibrant data. |  |  | Enter coefficients below graph.
Enter equation parameters below graph.
[10 empty] | Quadratic curve coefficients | [8 empty] | High curve coefficients:
[13 empty] | Intercept | X coef. | X^2 coef. |  |  |  |  | intcpt | X coef | X^2 coef
[13 empty] | -0.0032 | 1.2818 | 0.1054 |  |  |  |  | -0.0068 | 1.3523 | -0.0553
Show "Split" in spreadsheet.

## Apply Std Curve
|  |  | Sample Info |  |  |  | FASPac NO3-NO2 |  |  | Curve-fit results
SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | Peak# | Position | Identifier | Type | Raw Ht | Corr Ht | mg/L |  | mg/L | Diff | Curve-fitting parameters used:
2018-09-14 | 1 | a | 29 | 1:25 | 1 | Unknown | 0.043 | 0.04 | 0.048 |  | 0.046591 | 0.0014088 | Copy from Std Curve Construction page.
2018-09-14 | 1 | b | 30 | 1:26 | 2 | Unknown | 0.044 | 0.041 | 0.049 |  | 0.048281 | 0.0007186
2018-09-14 | 1 | c | 31 | 1:27 | 3 | Unknown | 0.042 | 0.039 | 0.047 |  | 0.044927 | 0.0020731 | Use these for linear or quadratic fit.
2018-09-14 | 2 | a | 32 | 1:28 | 4 | Unknown | 0.031 | 0.029 | 0.034 |  | 0.029707 | 0.0042933 | If linear, enter 0 for X^2 coefficient.
2018-09-14 | 2 | b | 33 | 1:29 | 5 | Unknown | 0.031 | 0.029 | 0.034 |  | 0.029707 | 0.0042933
2018-09-14 | 2 | c | 34 | 1:30 | 6 | Unknown | 0.032 | 0.029 | 0.035 |  | 0.029707 | 0.0052933 | Intercept: | X Coef. | X^2 Coef:
2018-09-14 | 3 | a | 35 | 1:31 | 7 | Unknown | 0.062 | 0.06 | 0.074 |  | 0.085311 | -0.011311
2018-09-14 | 3 | b | 36 | 1:32 | 8 | Unknown | 0.067 | 0.065 | 0.081 |  | 0.096608 | -0.015608
2018-09-14 | 3 | c | 37 | 1:33 | 9 | Unknown | 0.065 | 0.063 | 0.078 |  | 0.092012 | -0.014012
2018-09-14 | 4 | a | 38 | 1:34 | 10 | Unknown | 0.028 | 0.025 | 0.029 |  | 0.024343 | 0.0046569 | Use these for split curve:
2018-09-14 | 4 | b | 39 | 1:35 | 11 | Unknown | 0.029 | 0.026 | 0.031 |  | 0.025645 | 0.0053548 | Low Range:
2018-09-14 | 4 | c | 40 | 1:36 | 12 | Unknown | 0.027 | 0.025 | 0.029 |  | 0.024343 | 0.0046569 | Intercept: | X Coef. | X^2 Coef:
2018-09-14 | 5 | a | 41 | 1:37 | 13 | Unknown | 0.054 | 0.052 | 0.063 |  | 0.068581 | -0.0055812 | 0.0002 | 0.6423 | 12.937
2018-09-14 | 5 | b | 42 | 1:38 | 14 | Unknown | 0.042 | 0.04 | 0.048 |  | 0.046591 | 0.0014088
2018-09-14 | 5 | c | 43 | 1:39 | 15 | Unknown | 0.05 | 0.047 | 0.057 |  | 0.058966 | -0.0019659
2018-09-14 | 6 | a | 44 | 1:40 | 16 | Unknown | 0.106 | 0.103 | 0.13 |  | 0.1319 | -0.0019002 | High Range:
2018-09-14 | 6 | b | 45 | 1:41 | 17 | Unknown | 0.118 | 0.115 | 0.146 |  | 0.14798 | -0.0019832 | Intercept: | X Coef. | X^2 Coef:
2018-09-14 | 6 | c | 46 | 1:42 | 18 | Unknown | 0.114 | 0.111 | 0.14 |  | 0.14262 | -0.0026239 | -0.0068 | 1.3523 | -0.0553
2018-09-14 | 7 | a | 47 | 1:43 | 19 | Unknown | 0.108 | 0.105 | 0.132 |  | 0.13458 | -0.0025818
2018-09-14 | 7 | b | 48 | 1:44 | 20 | Unknown | 0.117 | 0.114 | 0.145 |  | 0.14664 | -0.0016435 | Procedure for split curve:
2018-09-14 | 7 | c | 49 | 1:45 | 21 | Unknown | 0.117 | 0.114 | 0.144 |  | 0.14664 | -0.0026435 | 1. Sort data on Row I: FASpac results, low to high.
2018-09-14 | 8 | a | 50 | 1:46 | 22 | Unknown | 0.11 | 0.107 | 0.135 |  | 0.13726 | -0.002263 | 2. Apply low range equation.
2018-09-14 | 8 | b | 51 | 1:47 | 23 | Unknown | 0.093 | 0.091 | 0.114 |  | 0.1158 | -0.0018014 | 3. Apply high range equation.
2018-09-14 | 8 | c | 52 | 1:48 | 24 | Unknown | 0.092 | 0.089 | 0.112 |  | 0.11312 | -0.0011167 | 4. Note here peak height where split occurs: |  |  | 0.01
2018-09-14 | 9 | a | 53 | 1:49 | 25 | Unknown | 0.009 | 0.006 | 0.005 |  | 0.0045195 | 0.00048047 | 4. Sort back based on Row D: Peak Number
2018-09-14 | 9 | b | 54 | 1:50 | 26 | Unknown | 0.008 | 0.005 | 0.004 |  | 0.0037349 | 0.00026508
2018-09-14 | 9 | c | 55 | 1:51 | 27 | Unknown | 0.007 | 0.005 | 0.003 |  | 0.0037349 | -0.00073492
2018-09-14 | 10 | a | 56 | 1:52 | 28 | Unknown | 0.005 | 0.003 | 0.001 |  | 0.0022433 | -0.0012433
2018-09-14 | 10 | b | 57 | 1:53 | 29 | Unknown | 0.005 | 0.003 | 0 |  | 0.0022433 | -0.0022433
2018-09-14 | 10 | c | 58 | 1:54 | 30 | Unknown | 0.007 | 0.004 | 0.002 |  | 0.0029762 | -0.00097619
2018-09-14 | 11 | a | 59 | 1:55 | 31 | Unknown | 0.006 | 0.003 | 0.001 |  | 0.0022433 | -0.0012433
2018-09-14 | 11 | b | 60 | 1:56 | 32 | Unknown | 0.007 | 0.004 | 0.002 |  | 0.0029762 | -0.00097619
2018-09-14 | 11 | c | 61 | 1:57 | 33 | Unknown | 0.006 | 0.003 | 0 |  | 0.0022433 | -0.0022433
2018-09-14 | 12 | a | 62 | 1:58 | 34 | Unknown | 0.011 | 0.008 | 0.007 |  | 0.0061664 | 0.00083363
2018-09-14 | 12 | b | 63 | 1:59 | 35 | Unknown | 0.011 | 0.008 | 0.007 |  | 0.0061664 | 0.00083363
2018-09-14 | 12 | c | 64 | 1:60 | 36 | Unknown | 0.011 | 0.007 | 0.006 |  | 0.00533 | 0.00066999
2018-09-14 | 13 | a | 65 | 2:1 | 37 | Unknown | 0.005 | 0.002 | -0.001 |  | 0.0015363 | -0.0025363
2018-09-14 | 13 | b | 66 | 2:2 | 38 | Unknown | 0.005 | 0.002 | -0.001 |  | 0.0015363 | -0.0025363
2018-09-14 | 13 | c | 67 | 2:3 | 39 | Unknown | 0.005 | 0.002 | -0.001 |  | 0.0015363 | -0.0025363
2018-09-27 | 1 | a | 68 | 2:4 | 40 | Unknown | 0.053 | 0.049 | 0.06 |  | 0.062734 | -0.0027344
2018-09-27 | 1 | b | 69 | 2:5 | 41 | Unknown | 0.059 | 0.055 | 0.067 |  | 0.074661 | -0.0076609
2018-09-27 | 1 | c | 70 | 2:6 | 42 | Unknown | 0.055 | 0.051 | 0.063 |  | 0.066606 | -0.0036064
2018-09-27 | 2 | a | 71 | 2:7 | 43 | Unknown | 0.033 | 0.029 | 0.034 |  | 0.029707 | 0.0042933
2018-09-27 | 2 | b |  | 2:8 | 44 | Unknown | 0.034 | 0.03 | 0.036 |  | 0.031112 | 0.0048877
2018-09-27 | 2 | c | 72 | 2:9 | 45 | Unknown | 0.032 | 0.028 | 0.033 |  | 0.028327 | 0.004673
2018-09-27 | 3 | a | 73 | 2:10 | 46 | Unknown | 0.089 | 0.085 | 0.106 |  | 0.10775 | -0.001746
2018-09-27 | 3 | b | 74 | 2:11 | 47 | Unknown | 0.081 | 0.076 | 0.095 |  | 0.095655 | -0.00065539
2018-09-27 | 3 | c | 75 | 2:12 | 48 | Unknown | 0.078 | 0.074 | 0.092 |  | 0.11857 | -0.026573
2018-09-27 | 4 | a | 76 | 2:13 | 49 | Unknown | 0.038 | 0.033 | 0.039 |  | 0.035484 | 0.0035157
2018-09-27 | 4 | b | 77 | 2:14 | 50 | Unknown | 0.038 | 0.033 | 0.04 |  | 0.035484 | 0.0045157
2018-09-27 | 4 | c | 78 | 2:15 | 51 | Unknown | 0.038 | 0.033 | 0.039 |  | 0.035484 | 0.0035157
2018-09-27 | 5 | a | 79 | 2:16 | 52 | Unknown | 0.05 | 0.045 | 0.055 |  | 0.055301 | -0.00030092
2018-09-27 | 5 | b | 80 | 2:17 | 53 | Unknown | 0.051 | 0.046 | 0.056 |  | 0.05712 | -0.0011205
2018-09-27 | 5 | c | 81 | 2:18 | 54 | Unknown | 0.057 | 0.052 | 0.064 |  | 0.068581 | -0.0045812
2018-09-27 | 6 | a | 82 | 2:19 | 55 | Unknown | 0.128 | 0.123 | 0.156 |  | 0.1587 | -0.0026963
2018-09-27 | 6 | b | 83 | 2:20 | 56 | Unknown | 0.132 | 0.127 | 0.161 |  | 0.16405 | -0.0030502
2018-09-27 | 6 | c | 84 | 2:21 | 57 | Unknown | 0.139 | 0.134 | 0.17 |  | 0.17342 | -0.0034152
2018-09-27 | 7 | a | 85 | 2:22 | 58 | Unknown | 0.121 | 0.116 | 0.147 |  | 0.14932 | -0.0023227
2018-09-27 | 7 | b | 86 | 2:23 | 59 | Unknown | 0.128 | 0.123 | 0.156 |  | 0.1587 | -0.0026963
2018-09-27 | 7 | c | 87 | 2:24 | 60 | Unknown | 0.11 | 0.106 | 0.133 |  | 0.13592 | -0.0029224
2018-09-27 | 8 | a | 88 | 2:25 | 61 | Unknown | 0.091 | 0.086 | 0.108 |  | 0.10909 | -0.0010888
2018-09-27 | 8 | b | 89 | 2:26 | 62 | Unknown | 0.089 | 0.084 | 0.106 |  | 0.1064 | -0.000403
2018-09-27 | 8 | c | 90 | 2:27 | 63 | Unknown | 0.094 | 0.09 | 0.112 |  | 0.11446 | -0.0024591
2018-09-27 | 9 | a | 91 | 2:28 | 64 | Unknown | 0.008 | 0.003 | 0.001 |  | 0.0022433 | -0.0012433
2018-09-27 | 9 | b | 92 | 2:29 | 65 | Unknown | 0.007 | 0.003 | 0.001 |  | 0.0022433 | -0.0012433
2018-09-27 | 9 | c | 93 | 2:30 | 66 | Unknown | 0.008 | 0.004 | 0.001 |  | 0.0029762 | -0.0019762
2018-09-27 | 10 | a | 94 | 2:31 | 67 | Unknown | 0.009 | 0.005 | 0.004 |  | 0.0037349 | 0.00026508
2018-09-27 | 10 | b | 95 | 2:32 | 68 | Unknown | 0.009 | 0.005 | 0.004 |  | 0.0037349 | 0.00026508
2018-09-27 | 10 | c | 96 | 2:33 | 69 | Unknown | 0.008 | 0.005 | 0.003 |  | 0.0037349 | -0.00073492
2018-09-27 | 11 | a | 97 | 2:34 | 70 | Unknown | 0.008 | 0.005 | 0.003 |  | 0.0037349 | -0.00073492
2018-09-27 | 11 | b | 98 | 2:35 | 71 | Unknown | 0.007 | 0.004 | 0.002 |  | 0.0029762 | -0.00097619
2018-09-27 | 11 | c | 99 | 2:36 | 72 | Unknown | 0.007 | 0.004 | 0.002 |  | 0.0029762 | -0.00097619
2018-09-27 | 12 | a | 100 | 2:37 | 73 | Unknown | 0.013 | 0.009 | 0.009 |  | 0.0070286 | 0.0019714
2018-09-27 | 12 | b | 101 | 2:38 | 74 | Unknown | 0.013 | 0.009 | 0.008 |  | 0.0070286 | 0.0009714
2018-09-27 | 12 | c | 102 | 2:39 | 75 | Unknown | 0.014 | 0.01 | 0.01 |  | 0.0079167 | 0.0020833
2018-09-27 | 13 | a | 103 | 2:40 | 76 | Unknown | 0.007 | 0.003 | 0.001 |  | 0.0022433 | -0.0012433
2018-09-27 | 13 | b | 104 | 2:41 | 77 | Unknown | 0.006 | 0.002 | 0 |  | 0.0015363 | -0.0015363
2018-09-27 | 13 | c | 105 | 2:42 | 78 | Unknown | 0.006 | 0.002 | 0 |  | 0.0015363 | -0.0015363
2018-10-11 | 1 | a | 106 | 2:43 | 79 | Unknown | 0.032 | 0.028 | 0.033 |  | 0.028327 | 0.004673
2018-10-11 | 1 | b | 107 | 2:44 | 80 | Unknown | 0.032 | 0.028 | 0.033 |  | 0.028327 | 0.004673
2018-10-11 | 1 | c | 108 | 2:45 | 81 | Unknown | 0.032 | 0.028 | 0.033 |  | 0.028327 | 0.004673
2018-10-11 | 2 | a | 109 | 2:46 | 82 | Unknown | 0.027 | 0.023 | 0.026 |  | 0.021817 | 0.0041834
2018-10-11 | 2 | b | 110 | 2:47 | 83 | Unknown | 0.026 | 0.022 | 0.025 |  | 0.020592 | 0.0044079
2018-10-11 | 2 | c | 111 | 2:48 | 84 | Unknown | 0.026 | 0.022 | 0.025 |  | 0.020592 | 0.0044079
2018-10-11 | 3 | a | 112 | 2:49 | 85 | Unknown | 0.061 | 0.057 | 0.07 |  | 0.078843 | -0.0088434
2018-10-11 | 3 | b | 113 | 2:50 | 86 | Unknown | 0.057 | 0.053 | 0.065 |  | 0.070582 | -0.0055819
2018-10-11 | 3 | c | 114 | 2:51 | 87 | Unknown | 0.066 | 0.062 | 0.076 |  | 0.089752 | -0.013752
2018-10-11 | 4 | a | 115 | 2:52 | 88 | Unknown | 0.046 | 0.042 | 0.05 |  | 0.049997 | 2.532e-06
2018-10-11 | 4 | b | 116 | 2:53 | 89 | Unknown | 0.046 | 0.042 | 0.051 |  | 0.049997 | 0.0010025
2018-10-11 | 4 | c | 117 | 2:54 | 90 | Unknown | 0.051 | 0.047 | 0.058 |  | 0.058966 | -0.00096593
2018-10-11 | 5 | a | 118 | 2:55 | 91 | Unknown | 0.063 | 0.059 | 0.072 |  | 0.083129 | -0.011129
2018-10-11 | 5 | b | 119 | 2:56 | 92 | Unknown | 0.065 | 0.061 | 0.075 |  | 0.087519 | -0.012519
2018-10-11 | 5 | c | 120 | 2:57 | 93 | Unknown | 0.063 | 0.059 | 0.073 |  | 0.083129 | -0.010129
2018-10-11 | 6 | a | 121 | 2:58 | 94 | Unknown | 0.102 | 0.098 | 0.123 |  | 0.12519 | -0.0021943
2018-10-11 | 6 | b | 122 | 2:59 | 95 | Unknown | 0.094 | 0.09 | 0.113 |  | 0.11446 | -0.0014591
2018-10-11 | 6 | c | 123 | 2:60 | 96 | Unknown | 0.088 | 0.083 | 0.104 |  | 0.10506 | -0.0010599
2018-10-11 | 7 | a | 124 | 3:1 | 97 | Unknown | 0.118 | 0.114 | 0.144 |  | 0.14664 | -0.0026435
2018-10-11 | 7 | b | 125 | 3:2 | 98 | Unknown | 0.1 | 0.096 | 0.12 |  | 0.12251 | -0.0025112
2018-10-11 | 7 | c | 126 | 3:3 | 99 | Unknown | 0.114 | 0.11 | 0.138 |  | 0.14128 | -0.0032839
2018-10-11 | 8 | a | 127 | 3:4 | 100 | Unknown | 0.11 | 0.105 | 0.133 |  | 0.13458 | -0.0015818
2018-10-11 | 8 | b | 128 | 3:5 | 101 | Unknown | 0.099 | 0.095 | 0.119 |  | 0.12117 | -0.0021694
2018-10-11 | 8 | c | 129 | 3:6 | 102 | Unknown | 0.095 | 0.091 | 0.114 |  | 0.1158 | -0.0018014
2018-10-11 | 9 | a | 130 | 3:7 | 103 | Unknown | 0.047 | 0.043 | 0.052 |  | 0.051739 | 0.00026059
2018-10-11 | 9 | b | 131 | 3:8 | 104 | Unknown | 0.049 | 0.045 | 0.054 |  | 0.055301 | -0.0013009
2018-10-11 | 9 | c | 132 | 3:9 | 105 | Unknown | 0.049 | 0.045 | 0.055 |  | 0.055301 | -0.00030092
2018-10-11 | 10 | a | 133 | 3:10 | 106 | Unknown | 0.021 | 0.017 | 0.019 |  | 0.014858 | 0.0041421
2018-10-11 | 10 | b | 134 | 3:11 | 107 | Unknown | 0.02 | 0.016 | 0.018 |  | 0.013789 | 0.0042113
2018-10-11 | 10 | c | 135 | 3:12 | 108 | Unknown | 0.017 | 0.014 | 0.014 |  | 0.011728 | 0.0022721
2018-10-11 | 11 | a | 136 | 3:13 | 109 | Unknown | 0.018 | 0.014 | 0.015 |  | 0.011728 | 0.0032721
2018-10-11 | 11 | b | 137 | 3:14 | 110 | Unknown | 0.013 | 0.009 | 0.009 |  | 0.0070286 | 0.0019714
2018-10-11 | 11 | c | 138 | 3:15 | 111 | Unknown | 0.017 | 0.013 | 0.014 |  | 0.010736 | 0.0032637
2018-10-11 | 12 | a | 139 | 3:16 | 112 | Unknown | 0.019 | 0.015 | 0.017 |  | 0.012745 | 0.0042547
2018-10-11 | 12 | b | 140 | 3:17 | 113 | Unknown | 0.018 | 0.014 | 0.015 |  | 0.011728 | 0.0032721
2018-10-11 | 12 | c | 141 | 3:18 | 114 | Unknown | 0.019 | 0.016 | 0.017 |  | 0.013789 | 0.0032113
2018-10-11 | 13 | a | 142 | 3:19 | 115 | Unknown | 0.009 | 0.005 | 0.004 |  | 0.0037349 | 0.00026508
2018-10-11 | 13 | b | 143 | 3:20 | 116 | Unknown | 0.009 | 0.005 | 0.003 |  | 0.0037349 | -0.00073492
2018-10-11 | 13 | c |  | 3:21 | 117 | Unknown | 0.008 | 0.005 | 0.003 |  | 0.0037349 | -0.00073492
2018-11-02 | 1 | a | 144 | 3:22 | 118 | Unknown | 0.027 | 0.023 | 0.026 |  | 0.021817 | 0.0041834
2018-11-02 | 1 | b | 145 | 3:23 | 119 | Unknown | 0.025 | 0.021 | 0.023 |  | 0.019394 | 0.0036065
2018-11-02 | 1 | c | 146 | 3:24 | 120 | Unknown | 0.026 | 0.023 | 0.026 |  | 0.021817 | 0.0041834
2018-11-02 | 2 | a | 147 | 3:25 | 121 | Unknown | 0.013 | 0.009 | 0.008 |  | 0.0070286 | 0.0009714
2018-11-02 | 2 | b | 148 | 3:26 | 122 | Unknown | 0.011 | 0.007 | 0.006 |  | 0.00533 | 0.00066999
2018-11-02 | 2 | c | 149 | 3:27 | 123 | Unknown | 0.012 | 0.008 | 0.007 |  | 0.0061664 | 0.00083363
2018-11-02 | 3 | a | 150 | 3:28 | 124 | Unknown | 0.051 | 0.047 | 0.057 |  | 0.058966 | -0.0019659
2018-11-02 | 3 | b | 151 | 3:29 | 125 | Unknown | 0.048 | 0.044 | 0.054 |  | 0.053507 | 0.00049277
2018-11-02 | 3 | c | 152 | 3:30 | 126 | Unknown | 0.051 | 0.047 | 0.057 |  | 0.058966 | -0.0019659
2018-11-02 | 4 | a | 153 | 3:31 | 127 | Unknown | 0.045 | 0.041 | 0.05 |  | 0.048281 | 0.0017186
2018-11-02 | 4 | b | 154 | 3:32 | 128 | Unknown | 0.045 | 0.04 | 0.049 |  | 0.046591 | 0.0024088
2018-11-02 | 4 | c | 155 | 3:33 | 129 | Unknown | 0.047 | 0.043 | 0.052 |  | 0.051739 | 0.00026059
2018-11-02 | 5 | a | 156 | 3:34 | 130 | Unknown | 0.06 | 0.056 | 0.069 |  | 0.076739 | -0.0077392
2018-11-02 | 5 | b | 157 | 3:35 | 131 | Unknown | 0.056 | 0.052 | 0.063 |  | 0.068581 | -0.0055812
2018-11-02 | 5 | c | 158 | 3:36 | 132 | Unknown | 0.057 | 0.053 | 0.065 |  | 0.070582 | -0.0055819
2018-11-02 | 6 | a | 159 | 3:37 | 133 | Unknown | 0.108 | 0.103 | 0.13 |  | 0.1319 | -0.0019002
2018-11-02 | 6 | b | 160 | 3:38 | 134 | Unknown | 0.102 | 0.098 | 0.123 |  | 0.12519 | -0.0021943
2018-11-02 | 6 | c | 161 | 3:39 | 135 | Unknown | 0.104 | 0.099 | 0.125 |  | 0.12654 | -0.0015357
2018-11-02 | 7 | a | 162 | 3:40 | 136 | Unknown | 0.098 | 0.093 | 0.117 |  | 0.11849 | -0.0014856
2018-11-02 | 7 | b | 163 | 3:41 | 137 | Unknown | 0.093 | 0.089 | 0.112 |  | 0.11312 | -0.0011167
2018-11-02 | 7 | c | 164 | 3:42 | 138 | Unknown | 0.105 | 0.1 | 0.127 |  | 0.12788 | -0.000877
2018-11-02 | 8 | a | 165 | 3:43 | 139 | Unknown | 0.097 | 0.093 | 0.117 |  | 0.11849 | -0.0014856
2018-11-02 | 8 | b | 166 | 3:44 | 140 | Unknown | 0.1 | 0.096 | 0.12 |  | 0.12251 | -0.0025112
2018-11-02 | 8 | c | 167 | 3:45 | 141 | Unknown | 0.093 | 0.089 | 0.111 |  | 0.11312 | -0.0021167
2018-11-02 | 9 | a | 168 | 3:46 | 142 | Unknown | 0.044 | 0.04 | 0.048 |  | 0.046591 | 0.0014088
2018-11-02 | 9 | b | 169 | 3:47 | 143 | Unknown | 0.042 | 0.038 | 0.045 |  | 0.043288 | 0.0017116
2018-11-02 | 9 | c | 170 | 3:48 | 144 | Unknown | 0.046 | 0.042 | 0.051 |  | 0.049997 | 0.0010025
2018-11-02 | 10 | a | 171 | 3:49 | 145 | Unknown | 0.051 | 0.047 | 0.057 |  | 0.058966 | -0.0019659
2018-11-02 | 10 | b | 172 | 3:50 | 146 | Unknown | 0.05 | 0.047 | 0.057 |  | 0.058966 | -0.0019659
2018-11-02 | 10 | c | 173 | 3:51 | 147 | Unknown | 0.052 | 0.049 | 0.059 |  | 0.062734 | -0.0037344
2018-11-02 | 11 | a | 174 | 3:52 | 148 | Unknown | 0.05 | 0.046 | 0.056 |  | 0.05712 | -0.0011205
2018-11-02 | 11 | b | 175 | 3:53 | 149 | Unknown | 0.05 | 0.046 | 0.056 |  | 0.05712 | -0.0011205
2018-11-02 | 11 | c | 176 | 3:54 | 150 | Unknown | 0.05 | 0.046 | 0.056 |  | 0.05712 | -0.0011205
2018-11-02 | 12 | a | 177 | 3:55 | 151 | Unknown | 0.044 | 0.04 | 0.049 |  | 0.046591 | 0.0024088
2018-11-02 | 12 | b | 178 | 3:56 | 152 | Unknown | 0.043 | 0.04 | 0.048 |  | 0.046591 | 0.0014088
2018-11-02 | 12 | c | 179 | 3:57 | 153 | Unknown | 0.043 | 0.039 | 0.048 |  | 0.044927 | 0.0030731
2018-11-02 | 13 | a | 180 | 3:58 | 154 | Unknown | 0.028 | 0.024 | 0.028 |  | 0.023067 | 0.0049331
2018-11-02 | 13 | b | 181 | 3:59 | 155 | Unknown | 0.028 | 0.024 | 0.028 |  | 0.023067 | 0.0049331
2018-11-02 | 13 | c | 182 | 3:60 | 156 | Unknown | 0.028 | 0.024 | 0.028 |  | 0.023067 | 0.0049331

## Final Results
SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | N-NO3 mg/L
2018-09-14 | 1 | a | 0.046591
2018-09-14 | 1 | b | 0.048281
2018-09-14 | 1 | c | 0.044927
2018-09-14 | 2 | a | 0.029707
2018-09-14 | 2 | b | 0.029707
2018-09-14 | 2 | c | 0.029707
2018-09-14 | 3 | a | 0.085311
2018-09-14 | 3 | b | 0.096608
2018-09-14 | 3 | c | 0.092012
2018-09-14 | 4 | a | 0.024343
2018-09-14 | 4 | b | 0.025645
2018-09-14 | 4 | c | 0.024343
2018-09-14 | 5 | a | 0.068581
2018-09-14 | 5 | b | 0.046591
2018-09-14 | 5 | c | 0.058966
2018-09-14 | 6 | a | 0.1319
2018-09-14 | 6 | b | 0.14798
2018-09-14 | 6 | c | 0.14262
2018-09-14 | 7 | a | 0.13458
2018-09-14 | 7 | b | 0.14664
2018-09-14 | 7 | c | 0.14664
2018-09-14 | 8 | a | 0.13726
2018-09-14 | 8 | b | 0.1158
2018-09-14 | 8 | c | 0.11312
2018-09-14 | 9 | a | 0.0045195
2018-09-14 | 9 | b | 0.0037349
2018-09-14 | 9 | c | 0.0037349
2018-09-14 | 10 | a | 0.0022433
2018-09-14 | 10 | b | 0.0022433
2018-09-14 | 10 | c | 0.0029762
2018-09-14 | 11 | a | 0.0022433
2018-09-14 | 11 | b | 0.0029762
2018-09-14 | 11 | c | 0.0022433
2018-09-14 | 12 | a | 0.0061664
2018-09-14 | 12 | b | 0.0061664
2018-09-14 | 12 | c | 0.00533
2018-09-14 | 13 | a | 0.0015363
2018-09-14 | 13 | b | 0.0015363
2018-09-14 | 13 | c | 0.0015363
2018-09-27 | 1 | a | 0.062734
2018-09-27 | 1 | b | 0.074661
2018-09-27 | 1 | c | 0.066606
2018-09-27 | 2 | a | 0.029707
2018-09-27 | 2 | b | 0.031112
2018-09-27 | 2 | c | 0.028327
2018-09-27 | 3 | a | 0.10775
2018-09-27 | 3 | b | 0.095655
2018-09-27 | 3 | c | 0.11857
2018-09-27 | 4 | a | 0.035484
2018-09-27 | 4 | b | 0.035484
2018-09-27 | 4 | c | 0.035484
2018-09-27 | 5 | a | 0.055301
2018-09-27 | 5 | b | 0.05712
2018-09-27 | 5 | c | 0.068581
2018-09-27 | 6 | a | 0.1587
2018-09-27 | 6 | b | 0.16405
2018-09-27 | 6 | c | 0.17342
2018-09-27 | 7 | a | 0.14932
2018-09-27 | 7 | b | 0.1587
2018-09-27 | 7 | c | 0.13592
2018-09-27 | 8 | a | 0.10909
2018-09-27 | 8 | b | 0.1064
2018-09-27 | 8 | c | 0.11446
2018-09-27 | 9 | a | 0.0022433
2018-09-27 | 9 | b | 0.0022433
2018-09-27 | 9 | c | 0.0029762
2018-09-27 | 10 | a | 0.0037349
2018-09-27 | 10 | b | 0.0037349
2018-09-27 | 10 | c | 0.0037349
2018-09-27 | 11 | a | 0.0037349
2018-09-27 | 11 | b | 0.0029762
2018-09-27 | 11 | c | 0.0029762
2018-09-27 | 12 | a | 0.0070286
2018-09-27 | 12 | b | 0.0070286
2018-09-27 | 12 | c | 0.0079167
2018-09-27 | 13 | a | 0.0022433
2018-09-27 | 13 | b | 0.0015363
2018-09-27 | 13 | c | 0.0015363
2018-10-11 | 1 | a | 0.028327
2018-10-11 | 1 | b | 0.028327
2018-10-11 | 1 | c | 0.028327
2018-10-11 | 2 | a | 0.021817
2018-10-11 | 2 | b | 0.020592
2018-10-11 | 2 | c | 0.020592
2018-10-11 | 3 | a | 0.078843
2018-10-11 | 3 | b | 0.070582
2018-10-11 | 3 | c | 0.089752
2018-10-11 | 4 | a | 0.049997
2018-10-11 | 4 | b | 0.049997
2018-10-11 | 4 | c | 0.058966
2018-10-11 | 5 | a | 0.083129
2018-10-11 | 5 | b | 0.087519
2018-10-11 | 5 | c | 0.083129
2018-10-11 | 6 | a | 0.12519
2018-10-11 | 6 | b | 0.11446
2018-10-11 | 6 | c | 0.10506
2018-10-11 | 7 | a | 0.14664
2018-10-11 | 7 | b | 0.12251
2018-10-11 | 7 | c | 0.14128
2018-10-11 | 8 | a | 0.13458
2018-10-11 | 8 | b | 0.12117
2018-10-11 | 8 | c | 0.1158
2018-10-11 | 9 | a | 0.051739
2018-10-11 | 9 | b | 0.055301
2018-10-11 | 9 | c | 0.055301
2018-10-11 | 10 | a | 0.014858
2018-10-11 | 10 | b | 0.013789
2018-10-11 | 10 | c | 0.011728
2018-10-11 | 11 | a | 0.011728
2018-10-11 | 11 | b | 0.0070286
2018-10-11 | 11 | c | 0.010736
2018-10-11 | 12 | a | 0.012745
2018-10-11 | 12 | b | 0.011728
2018-10-11 | 12 | c | 0.013789
2018-10-11 | 13 | a | 0.0037349
2018-10-11 | 13 | b | 0.0037349
2018-10-11 | 13 | c | 0.0037349
2018-11-02 | 1 | a | 0.021817
2018-11-02 | 1 | b | 0.019394
2018-11-02 | 1 | c | 0.021817
2018-11-02 | 2 | a | 0.0070286
2018-11-02 | 2 | b | 0.00533
2018-11-02 | 2 | c | 0.0061664
2018-11-02 | 3 | a | 0.058966
2018-11-02 | 3 | b | 0.053507
2018-11-02 | 3 | c | 0.058966
2018-11-02 | 4 | a | 0.048281
2018-11-02 | 4 | b | 0.046591
2018-11-02 | 4 | c | 0.051739
2018-11-02 | 5 | a | 0.076739
2018-11-02 | 5 | b | 0.068581
2018-11-02 | 5 | c | 0.070582
2018-11-02 | 6 | a | 0.1319
2018-11-02 | 6 | b | 0.12519
2018-11-02 | 6 | c | 0.12654
2018-11-02 | 7 | a | 0.11849
2018-11-02 | 7 | b | 0.11312
2018-11-02 | 7 | c | 0.12788
2018-11-02 | 8 | a | 0.11849
2018-11-02 | 8 | b | 0.12251
2018-11-02 | 8 | c | 0.11312
2018-11-02 | 9 | a | 0.046591
2018-11-02 | 9 | b | 0.043288
2018-11-02 | 9 | c | 0.049997
2018-11-02 | 10 | a | 0.058966
2018-11-02 | 10 | b | 0.058966
2018-11-02 | 10 | c | 0.062734
2018-11-02 | 11 | a | 0.05712
2018-11-02 | 11 | b | 0.05712
2018-11-02 | 11 | c | 0.05712
2018-11-02 | 12 | a | 0.046591
2018-11-02 | 12 | b | 0.046591
2018-11-02 | 12 | c | 0.044927
2018-11-02 | 13 | a | 0.023067
2018-11-02 | 13 | b | 0.023067
2018-11-02 | 13 | c | 0.023067