## Results and Sample ID
Run Name: LTREB_NH4-PO4_AP2_2019_09_26
Run Date: 9/26/2019
Operator: Valett Lab
|  |  |  | Sample Info |  |  |  | Ammonium |  |  | Phosphate |  |  | Calibrant | Conc., mg/L
Row | SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | Peak# | Position | Identifier | Type | Raw Ht | Cor Ht | ppm | Raw Ht | Cor Ht | ppm | C1 | 0.001
|  |  |  | 1 | 1:1 | SYNC | SYNC | 0.039 | 0.0384 | 0.2037 | 0.0494 | 0.0499 | 0.2045 | C2 | 0.005
|  |  |  | 2 | 1:2 | CO | Carry over | 0.0011 | 0.0005 | 0.0051 | -0.0003 | 0.0002 | 0.001 | C3 | 0.01
|  |  |  | 3 | 1:3 | W | Wash | 0.0006 | 0 | 0.0024 | -0.0005 | 0 | 0.0003 | C4 | 0.05
|  |  |  | 4 | 1:3 | W | Wash | 0.0006 | 0 | 0.0024 | -0.0005 | 0 | 0.0003 | C5 | 0.1
|  |  |  | 5 | 1:4 | C1 | Calibrant | 0.0006 | 0 | 0.0024 | -0.0002 | 0.0002 | 0.0012 | C6 | 0.2
|  |  |  | 6 | 1:4 | C1 | Calibrant | 0.0007 | 0 | 0.0026 | -0.0003 | 0.0002 | 0.001
|  |  |  | 7 | 1:4 | C1 | Calibrant | 0.0007 | 0 | 0.0026 | -0.0003 | 0.0002 | 0.0011
|  |  |  | 8 | 1:5 | C2 | Calibrant | 0.0011 | 0.0004 | 0.0047 | 0.0005 | 0.0009 | 0.0042
|  |  |  | 9 | 1:5 | C2 | Calibrant | 0.0011 | 0.0005 | 0.0048 | 0.0006 | 0.001 | 0.0043
|  |  |  | 10 | 1:5 | C2 | Calibrant | 0.0011 | 0.0004 | 0.0045 | 0.0006 | 0.001 | 0.0043
|  |  |  | 11 | 1:6 | C3 | Calibrant | 0.0024 | 0.0017 | 0.0113 | 0.0018 | 0.0022 | 0.0093
|  |  |  | 12 | 1:6 | C3 | Calibrant | 0.0025 | 0.0017 | 0.0116 | 0.0019 | 0.0022 | 0.0095
|  |  |  | 13 | 1:6 | C3 | Calibrant | 0.0023 | 0.0015 | 0.0104 | 0.002 | 0.0023 | 0.0099
|  |  |  | 14 | 1:7 | C4 | Calibrant | 0.009 | 0.0083 | 0.0457 | 0.0112 | 0.0115 | 0.0476
|  |  |  | 15 | 1:7 | C4 | Calibrant | 0.0087 | 0.0079 | 0.0436 | 0.0114 | 0.0117 | 0.0482
|  |  |  | 16 | 1:7 | C4 | Calibrant | 0.0089 | 0.008 | 0.0443 | 0.0113 | 0.0116 | 0.0478
|  |  |  | 17 | 1:8 | C5 | Calibrant | 0.0192 | 0.0183 | 0.0983 | 0.025 | 0.0253 | 0.1041
|  |  |  | 18 | 1:8 | C5 | Calibrant | 0.0211 | 0.0201 | 0.1076 | 0.0254 | 0.0257 | 0.1054
|  |  |  | 19 | 1:8 | C5 | Calibrant | 0.0207 | 0.0196 | 0.105 | 0.0256 | 0.0258 | 0.1061
|  |  |  | 20 | 1:9 | C6 | Calibrant | 0.039 | 0.0379 | 0.201 | 0.0478 | 0.0481 | 0.197
|  |  |  | 21 | 1:9 | C6 | Calibrant | 0.039 | 0.0376 | 0.1996 | 0.0481 | 0.0483 | 0.1978
|  |  |  | 22 | 1:9 | C6 | Calibrant | 0.0387 | 0.0373 | 0.1979 | 0.0484 | 0.0486 | 0.1991
|  |  |  | 23 | 1:3 | W | Wash | 0.0009 | 0 | 0.0024 | -0.0003 | 0 | 0.0003
|  |  |  | 24 | 1:10 | CC1 | Unknown | 0.009 | 0.0082 | 0.0452 | 0.0112 | 0.0115 | 0.0473
|  |  |  | 25 | 1:11 | CC2 | Unknown | 0.0199 | 0.019 | 0.1021 | 0.0249 | 0.0252 | 0.1034
| 2019-06-10 | 1 | a | 26 | 1:25 | 1 | Unknown | 0.002 | 0.0011 | 0.008 | 0.0039 | 0.0042 | 0.0174
| 2019-06-10 | 1 | b | 27 | 1:26 | 2 | Unknown | 0.002 | 0.0014 | 0.0096 | 0.0039 | 0.0042 | 0.0176
| 2019-06-10 | 1 | c | 28 | 1:27 | 3 | Unknown | 0.0015 | 0.001 | 0.0075 | 0.0044 | 0.0048 | 0.02
| 2019-06-10 | 2 | a | 29 | 1:28 | 4 | Unknown | 0.0023 | 0.0019 | 0.0122 | 0.0029 | 0.0033 | 0.0137
| 2019-06-10 | 2 | b | 30 | 1:29 | 5 | Unknown | 0.0018 | 0.0014 | 0.0096 | 0.0034 | 0.0038 | 0.0158
| 2019-06-10 | 2 | c | 31 | 1:30 | 6 | Unknown | 0.0012 | 0.0009 | 0.007 | 0.0047 | 0.0052 | 0.0215
| 2019-06-10 | 3 | a | 32 | 1:31 | 7 | Unknown | 0.0012 | 0.001 | 0.0075 | 0.0036 | 0.004 | 0.0168
| 2019-06-10 | 3 | b | 33 | 1:32 | 8 | Unknown | 0.0015 | 0.0013 | 0.0091 | 0.0032 | 0.0037 | 0.0154
| 2019-06-10 | 3 | c | 35 | 1:33 | 9 | Unknown | 0.0015 | 0.0013 | 0.0095 | 0.0039 | 0.0044 | 0.0184
| 2019-06-10 | 4 | a | 36 | 1:34 | 10 | Unknown | 0.0015 | 0.0014 | 0.0096 | 0.0018 | 0.0022 | 0.0095
| 2019-06-10 | 4 | b | 37 | 1:35 | 11 | Unknown | 0.0021 | 0.002 | 0.0128 | 0.0049 | 0.0053 | 0.0222
| 2019-06-10 | 4 | c | 38 | 1:36 | 12 | Unknown | 0.0014 | 0.0013 | 0.009 | 0.0041 | 0.0045 | 0.0189
| 2019-06-10 | 5 | a | 39 | 1:37 | 13 | Unknown | 0.0014 | 0.0013 | 0.0093 | 0.0035 | 0.004 | 0.0165
| 2019-06-10 | 5 | b | 40 | 1:38 | 14 | Unknown | 0.0015 | 0.0013 | 0.0095 | 0.0025 | 0.003 | 0.0125
| 2019-06-10 | 5 | c | 41 | 1:39 | 15 | Unknown | 0.0022 | 0.0021 | 0.0132 | 0.0041 | 0.0045 | 0.0189
| 2019-06-10 | 6 | a | 42 | 1:40 | 16 | Unknown | 0.0009 | 0.0008 | 0.0064 | 0.0021 | 0.0026 | 0.0109
| 2019-06-10 | 6 | b | 43 | 1:41 | 17 | Unknown | 0.0009 | 0.0008 | 0.0065 | 0.0036 | 0.0041 | 0.0172
| 2019-06-10 | 6 | c | 44 | 1:42 | 18 | Unknown | 0.0015 | 0.0014 | 0.01 | 0.0043 | 0.0047 | 0.0197
| 2019-06-10 | 7 | a | 46 | 1:43 | 19 | Unknown | 0.002 | 0.0019 | 0.0124 | 0.0048 | 0.0052 | 0.0218
| 2019-06-10 | 7 | b | 47 | 1:44 | 20 | Unknown | 0.0021 | 0.0019 | 0.0124 | 0.0025 | 0.003 | 0.0126
| 2019-06-10 | 7 | c | 48 | 1:45 | 21 | Unknown | 0.0012 | 0.001 | 0.0076 | 0.0037 | 0.0041 | 0.0173
| 2019-06-10 | 8 | a | 49 | 1:46 | 22 | Unknown | 0.0009 | 0.0007 | 0.0062 | 0.0047 | 0.0052 | 0.0216
| 2019-06-10 | 8 | b | 50 | 1:47 | 23 | Unknown | 0.0017 | 0.0015 | 0.0104 | 0.0049 | 0.0054 | 0.0223
| 2019-06-10 | 8 | c | 51 | 1:48 | 24 | Unknown | 0.0016 | 0.0014 | 0.0096 | 0.0044 | 0.0048 | 0.0201
| 2019-06-10 | 9 | a | 52 | 1:49 | 25 | Unknown | 0.0011 | 0.0008 | 0.0067 | 0.0024 | 0.0029 | 0.0121
| 2019-06-10 | 9 | b | 53 | 1:50 | 26 | Unknown | 0.0014 | 0.0012 | 0.0086 | 0.0042 | 0.0047 | 0.0197
| 2019-06-10 | 9 | c | 54 | 1:51 | 27 | Unknown | 0.0009 | 0.0006 | 0.0054 | 0.0038 | 0.0043 | 0.0179
| 2019-06-10 | 10 | a | 55 | 1:52 | 28 | Unknown | 0.0005 | 0.0002 | 0.0035 | 0.0034 | 0.0039 | 0.0163
| 2019-06-10 | 10 | b | 57 | 1:53 | 29 | Unknown | 0.001 | 0.0008 | 0.0064 | 0.0055 | 0.006 | 0.0248
| 2019-06-10 | 10 | c | 58 | 1:54 | 30 | Unknown | 0.0007 | 0.0004 | 0.0046 | 0.0042 | 0.0047 | 0.0195
| 2019-06-10 | 11 | a | 59 | 1:55 | 31 | Unknown | 0.0012 | 0.001 | 0.0077 | 0.0032 | 0.0037 | 0.0157
| 2019-06-10 | 11 | b | 60 | 1:56 | 32 | Unknown | 0.0012 | 0.001 | 0.0075 | 0.0026 | 0.0032 | 0.0133
| 2019-06-10 | 11 | c | 61 | 1:57 | 33 | Unknown | 0.0007 | 0.0005 | 0.0049 | 0.0022 | 0.0028 | 0.0117
| 2019-06-10 | 12 | a | 62 | 1:58 | 34 | Unknown | 0.0013 | 0.0011 | 0.0083 | 0.0018 | 0.0024 | 0.0102
| 2019-06-10 | 12 | b | 63 | 1:59 | 35 | Unknown | 0.0007 | 0.0006 | 0.0055 | 0.0034 | 0.004 | 0.0165
| 2019-06-10 | 12 | c | 64 | 1:60 | 36 | Unknown | 0.0011 | 0.001 | 0.0077 | 0.0023 | 0.0029 | 0.0121
| 2019-06-10 | 13 | a | 65 | 2:1 | 37 | Unknown | 0.0012 | 0.0011 | 0.0084 | 0.0009 | 0.0015 | 0.0064
| 2019-06-10 | 13 | b | 66 | 2:2 | 38 | Unknown | 0.0007 | 0.0006 | 0.0057 | 0.0017 | 0.0023 | 0.0098
| 2019-06-10 | 13 | c | 68 | 2:3 | 39 | Unknown | 0.0008 | 0.0008 | 0.0065 | 0.0011 | 0.0018 | 0.0076
| 2019-06-25 | 1 | a | 69 | 2:4 | 40 | Unknown | 0.0012 | 0.0012 | 0.0087 | 0.0042 | 0.0048 | 0.02
| 2019-06-25 | 1 | b | 70 | 2:5 | 41 | Unknown | 0.0008 | 0.0009 | 0.007 | 0.0039 | 0.0045 | 0.0186
| 2019-06-25 | 1 | c | 71 | 2:6 | 42 | Unknown | 0.001 | 0.001 | 0.0078 | 0.0044 | 0.005 | 0.021
| 2019-06-25 | 2 | a | 72 | 2:7 | 43 | Unknown | 0.0014 | 0.0015 | 0.01 | 0.0027 | 0.0033 | 0.014
| 2019-06-25 | 2 | b | 73 | 2:8 | 44 | Unknown | 0.8956 | 0.8956 | 4.6968 | 1.3846 | 1.3852 | 5.6692
| 2019-06-25 | 2 | c | 74 | 2:9 | 45 | Unknown | 0.0013 | -0.0105 | -0.0527 | 0.0037 | 0 | 0.0004
| 2019-06-25 | 3 | a | 75 | 2:10 | 46 | Unknown | 0.0008 | 0.001 | 0.0076 | 0.0045 | 0.0051 | 0.0211
| 2019-06-25 | 3 | b | 76 | 2:11 | 47 | Unknown | 0.001 | 0.001 | 0.0079 | 0.0041 | 0.0047 | 0.0196
| 2019-06-25 | 3 | c | 77 | 2:12 | 48 | Unknown | 0.0015 | 0.0016 | 0.0108 | 0.0041 | 0.0047 | 0.0194
| 2019-06-25 | 4 | a | 79 | 2:13 | 49 | Unknown | 0.0017 | 0.0017 | 0.0115 | 0.0034 | 0.004 | 0.0167
| 2019-06-25 | 4 | b | 80 | 2:14 | 50 | Unknown | 0.0004 | 0.0004 | 0.0044 | -0.0006 | -0.0001 | 0.0001
| 2019-06-25 | 4 | c | 83 | 2:15 | 51 | Unknown | 0.0033 | 0.003 | 0.0183 | 0.0008 | 0.0014 | 0.0059
| 2019-06-25 | 5 | a | 84 | 2:16 | 52 | Unknown | 0.001 | 0.0009 | 0.0071 | 0.0022 | 0.0028 | 0.0119
| 2019-06-25 | 5 | b | 85 | 2:17 | 53 | Unknown | 0.0012 | 0.0011 | 0.008 | 0.0026 | 0.0033 | 0.0138
| 2019-06-25 | 5 | c | 86 | 2:18 | 54 | Unknown | 0.0009 | 0.0008 | 0.0065 | 0.0025 | 0.0031 | 0.0131
| 2019-06-25 | 6 | a | 87 | 2:19 | 55 | Unknown | 0.0006 | 0.0005 | 0.005 | 0.0013 | 0.002 | 0.0085
| 2019-06-25 | 6 | b | 88 | 2:20 | 56 | Unknown | 0.0006 | 0.0005 | 0.0051 | 0.0009 | 0.0016 | 0.0068
| 2019-06-25 | 6 | c | 90 | 2:21 | 57 | Unknown | 0.0013 | 0.0012 | 0.0085 | 0.0006 | 0.0013 | 0.0056
| 2019-06-25 | 7 | a | 91 | 2:22 | 58 | Unknown | 0.0009 | 0.0008 | 0.0064 | 0.0009 | 0.0016 | 0.007
| 2019-06-25 | 7 | b | 92 | 2:23 | 59 | Unknown | 0.9384 | 0.9382 | 4.9202 | 1.6165 | 1.6172 | 6.6186
| 2019-06-25 | 7 | c | 93 | 2:24 | 60 | Unknown | 0.0017 | -0.0108 | -0.0544 | 0.0009 | -0.0034 | -0.0135
| 2019-06-25 | 8 | a | 94 | 2:25 | 61 | Unknown | 0.0013 | 0.0013 | 0.0092 | 0.0014 | 0.0021 | 0.0091
| 2019-06-25 | 8 | b | 95 | 2:26 | 62 | Unknown | 0.9161 | 0.9159 | 4.8031 | 1.5517 | 1.5523 | 6.353
| 2019-06-25 | 8 | c | 96 | 2:27 | 63 | Unknown | 0.0005 | -0.0117 | -0.0591 | -0.0007 | -0.0048 | -0.0194
| 2019-06-25 | 9 | a | 97 | 2:28 | 64 | Unknown | 0.0005 | 0.0005 | 0.005 | 0.0019 | 0.0026 | 0.011
| 2019-06-25 | 9 | b | 98 | 2:29 | 65 | Unknown | 0.0005 | 0.0004 | 0.0044 | -0.0005 | 0.0001 | 0.0008
| 2019-06-25 | 9 | c | 99 | 2:30 | 66 | Unknown | 0.0006 | 0.0005 | 0.0049 | 0.0023 | 0.003 | 0.0125
| 2019-06-25 | 10 | a | 101 | 2:31 | 67 | Unknown | 0.0003 | 0.0001 | 0.0031 | 0.004 | 0.0047 | 0.0194
| 2019-06-25 | 10 | b | 102 | 2:32 | 68 | Unknown | 0.0002 | 0.0001 | 0.0027 | 0.0037 | 0.0044 | 0.0182
| 2019-06-25 | 10 | c | 103 | 2:33 | 69 | Unknown | 0.0007 | 0.0005 | 0.0052 | 0.0031 | 0.0037 | 0.0157
| 2019-06-25 | 11 | a | 104 | 2:34 | 70 | Unknown | 0.0006 | 0.0004 | 0.0047 | 0.0026 | 0.0032 | 0.0135
| 2019-06-25 | 11 | b | 105 | 2:35 | 71 | Unknown | 0.0078 | 0.0076 | 0.0423 | 0.0018 | 0.0024 | 0.0103
| 2019-06-25 | 11 | c | 106 | 2:36 | 72 | Unknown | 0.0008 | 0.0006 | 0.0056 | 0.0019 | 0.0024 | 0.0103
| 2019-06-25 | 12 | a | 107 | 2:37 | 73 | Unknown | 0.0019 | 0.0017 | 0.0115 | 0.0016 | 0.0021 | 0.0091
| 2019-06-25 | 12 | b | 108 | 2:38 | 74 | Unknown | 0.9226 | 0.9225 | 4.8376 | 1.5634 | 1.564 | 6.4006
| 2019-06-25 | 12 | c | 109 | 2:39 | 75 | Unknown | 0.0012 | -0.0111 | -0.0559 | 0.0018 | -0.0024 | -0.0096
| 2019-06-25 | 13 | a | 110 | 2:40 | 76 | Unknown | 0.0004 | 0.0004 | 0.0045 | 0.0007 | 0.0013 | 0.0055
| 2019-06-25 | 13 | b | 112 | 2:41 | 77 | Unknown | 0.0027 | 0.0025 | 0.0156 | -0.0006 | 0 | 0.0003
| 2019-06-25 | 13 | c | 113 | 2:42 | 78 | Unknown | 0.0008 | 0.0006 | 0.0054 | 0.0006 | 0.0012 | 0.0052
| 2019-07-09 | 1 | a | 114 | 2:43 | 79 | Unknown | 0.002 | 0.0018 | 0.012 | 0.0047 | 0.0052 | 0.0217
| 2019-07-09 | 1 | b | 115 | 2:44 | 80 | Unknown | 0.0013 | 0.001 | 0.0079 | 0.0049 | 0.0054 | 0.0225
| 2019-07-09 | 1 | c | 116 | 2:45 | 81 | Unknown | 0.0037 | 0.0035 | 0.0206 | 0.0057 | 0.0063 | 0.0259
| 2019-07-09 | 2 | a | 119 | 2:46 | 82 | Unknown | 0.0019 | 0.0013 | 0.0094 | 0.0061 | 0.0066 | 0.0273
| 2019-07-09 | 2 | b | 120 | 2:47 | 83 | Unknown | 0.0016 | 0.0012 | 0.0088 | 0.0056 | 0.0062 | 0.0258
| 2019-07-09 | 2 | c | 121 | 2:48 | 84 | Unknown | 0.0018 | 0.0014 | 0.01 | 0.0049 | 0.0054 | 0.0226
| 2019-07-09 | 3 | a | 123 | 2:49 | 85 | Unknown | 0.0015 | 0.0012 | 0.0085 | 0.0051 | 0.0057 | 0.0237
| 2019-07-09 | 3 | b | 124 | 2:50 | 86 | Unknown | 0.0014 | 0.001 | 0.0079 | 0.0042 | 0.0049 | 0.0203
| 2019-07-09 | 3 | c | 125 | 2:51 | 87 | Unknown | 0.0016 | 0.0013 | 0.0092 | 0.0039 | 0.0046 | 0.019
| 2019-07-09 | 4 | a | 126 | 2:52 | 88 | Unknown | 0.0016 | 0.0012 | 0.0089 | 0.0039 | 0.0046 | 0.0191
| 2019-07-09 | 4 | b | 127 | 2:53 | 89 | Unknown | 0.0011 | 0.0008 | 0.0066 | 0.003 | 0.0037 | 0.0154
| 2019-07-09 | 4 | c | 128 | 2:54 | 90 | Unknown | 0.0011 | 0.0009 | 0.0069 | 0.0037 | 0.0043 | 0.0181
| 2019-07-09 | 5 | a | 129 | 2:55 | 91 | Unknown | 0.0013 | 0.001 | 0.0079 | 0.003 | 0.0037 | 0.0154
| 2019-07-09 | 5 | b | 130 | 2:56 | 92 | Unknown | 0.0012 | 0.0009 | 0.0073 | 0.0028 | 0.0035 | 0.0147
| 2019-07-09 | 5 | c | 131 | 2:57 | 93 | Unknown | 0.0015 | 0.0013 | 0.0091 | 0.0022 | 0.0029 | 0.012
| 2019-07-09 | 6 | a | 132 | 2:58 | 94 | Unknown | 0.0011 | 0.0008 | 0.0068 | 0.0019 | 0.0026 | 0.0111
| 2019-07-09 | 6 | b | 134 | 2:59 | 95 | Unknown | 0.0011 | 0.0008 | 0.0068 | 0.0016 | 0.0023 | 0.0096
| 2019-07-09 | 6 | c | 135 | 2:60 | 96 | Unknown | 0.0016 | 0.0014 | 0.0098 | 0.0014 | 0.002 | 0.0087
| 2019-07-09 | 7 | a | 136 | 3:1 | 97 | Unknown | 0.0009 | 0.0006 | 0.0056 | 0.0015 | 0.0022 | 0.0094
| 2019-07-09 | 7 | b | 137 | 3:2 | 98 | Unknown | 0.0008 | 0.0006 | 0.0054 | 0.0022 | 0.0029 | 0.0122
| 2019-07-09 | 7 | c | 138 | 3:3 | 99 | Unknown | 0.0024 | 0.0021 | 0.0135 | 0.0018 | 0.0024 | 0.0102
| 2019-07-09 | 8 | a | 139 | 3:4 | 100 | Unknown | 0.0007 | 0.0004 | 0.0044 | 0.0035 | 0.0042 | 0.0175
| 2019-07-09 | 8 | b | 140 | 3:5 | 101 | Unknown | 0.0007 | 0.0004 | 0.0045 | 0.0029 | 0.0035 | 0.0146
| 2019-07-09 | 8 | c | 141 | 3:6 | 102 | Unknown | 0.0008 | 0.0005 | 0.005 | 0.0036 | 0.0042 | 0.0175
| 2019-07-09 | 9 | a | 142 | 3:7 | 103 | Unknown | 0.0009 | 0.0006 | 0.0056 | 0.0031 | 0.0037 | 0.0155
| 2019-07-09 | 9 | b | 143 | 3:8 | 104 | Unknown | 0.0009 | 0.0005 | 0.0053 | 0.0034 | 0.004 | 0.0166
| 2019-07-09 | 9 | c | 145 | 3:9 | 105 | Unknown | 0.001 | 0.0007 | 0.0061 | 0.0025 | 0.0031 | 0.0129
| 2019-07-09 | 10 | a | 146 | 3:10 | 106 | Unknown | 0.0029 | 0.0025 | 0.0156 | 0.0041 | 0.0047 | 0.0197
| 2019-07-09 | 10 | b | 147 | 3:11 | 107 | Unknown | 0.0012 | 0.0009 | 0.0071 | 0.0036 | 0.0042 | 0.0176
| 2019-07-09 | 10 | c | 148 | 3:12 | 108 | Unknown | 0.001 | 0.0007 | 0.0062 | 0.0038 | 0.0044 | 0.0184
| 2019-07-09 | 11 | a | 149 | 3:13 | 109 | Unknown | 0.0009 | 0.0006 | 0.0056 | 0.003 | 0.0036 | 0.0152
| 2019-07-09 | 11 | b | 150 | 3:14 | 110 | Unknown | 0.0011 | 0.0008 | 0.0066 | 0.0026 | 0.0032 | 0.0134
| 2019-07-09 | 11 | c | 151 | 3:15 | 111 | Unknown | 0.0011 | 0.0007 | 0.0063 | 0.003 | 0.0036 | 0.015
| 2019-07-09 | 12 | a | 152 | 3:16 | 112 | Unknown | 0.0013 | 0.001 | 0.0078 | 0.0013 | 0.0019 | 0.0082
| 2019-07-09 | 12 | b | 153 | 3:17 | 113 | Unknown | 0.0012 | 0.0009 | 0.0071 | 0.001 | 0.0016 | 0.007
| 2019-07-09 | 12 | c | 154 | 3:18 | 114 | Unknown | 0.0011 | 0.0008 | 0.0067 | 0.0019 | 0.0025 | 0.0106
| 2019-07-09 | 13 | a | 156 | 3:19 | 115 | Unknown | 0.0016 | 0.0013 | 0.0094 | 0.0004 | 0.001 | 0.0043
| 2019-07-09 | 13 | b | 157 | 3:20 | 116 | Unknown | 0.0011 | 0.0008 | 0.0067 | 0.0002 | 0.0007 | 0.0034
| 2019-07-09 | 13 | c | 158 | 3:21 | 117 | Unknown | 0.0019 | 0.0017 | 0.0112 | 0.0003 | 0.0008 | 0.0037

## Std Curve Construction
1. Copy in Calibrant Information | [5 empty] | NH3 SECTION | [6 empty] | PO4 SECTION
[18 empty] | 2. Run Linear Regression | [8 empty] | 4. If quadratic exceeds 10%, run split-curve procedure.
Use X=Peak height; Y=Conc)
| Calibrant |  |  | Linear Fit |  | Quadratic |  | Split Curve |  |  | Linear Fit |  | Quadratic |  | Split Curve
Calibrant | Conc. | NH3 Calibrant | PO4 Calibrant | Std Conc |  | Std Conc |  | Std Conc |  |  | Std Conc |  | Std Conc |  | Std Conc | [11 empty] | Plot low range calibrant data: |  |  | Enter coefficients below graph.
Id | (mg/L N) | peak height (x) | peak height (x) | (y-hat) | % error | (y-hat) | % error | (y-hat) | % error |  | (y-hat) | % error | (y-hat) | % error | (y-hat) | % error |  |  | All data - Linear Fit
C1 | 0.001 | 0 | 0.0002 | 0.0025686 | 156.8607 | 0.0026 | 160 | 0.0017 | 70 |  | 0.0012738 | 27.3838 | 0.0023575 | 135.7475 | 0.0011863 | 18.6271 |  | SUMMARY OUTPUT
C1 | 0.001 | 0 | 0.0002 | 0.0025686 | 156.8607 | 0.0026 | 160 | 0.0017 | 70 |  | 0.0012738 | 27.3838 | 0.0023575 | 135.7475 | 0.0011863 | 18.6271
C1 | 0.001 | 0 | 0.0002 | 0.0025686 | 156.8607 | 0.0026 | 160 | 0.0017 | 70 |  | 0.0012738 | 27.3838 | 0.0023575 | 135.7475 | 0.0011863 | 18.6271 |  | Regression Statistics
C2 | 0.005 | 0.0004 | 0.0009 | 0.0046596 | -6.8082 | 0.0046804 | -6.3928 | 0.0038048 | -23.9038 |  | 0.0041331 | -17.3386 | 0.0050127 | 0.25317 | 0.0042799 | -14.4024 |  | Multiple R | 0.99939
C2 | 0.005 | 0.0005 | 0.001 | 0.0051823 | 3.6467 | 0.0052005 | 4.0097 | 0.0043357 | -13.2857 |  | 0.0045415 | -9.1694 | 0.0053925 | 7.8496 | 0.0047208 | -5.5847 |  | R Square | 0.99879
C2 | 0.005 | 0.0004 | 0.001 | 0.0046596 | -6.8082 | 0.0046804 | -6.3928 | 0.0038048 | -23.9038 |  | 0.0045415 | -9.1694 | 0.0053925 | 7.8496 | 0.0047208 | -5.5847 |  | Adjusted R Square | 0.99871
C3 | 0.01 | 0.0017 | 0.0022 | 0.011455 | 14.5529 | 0.011443 | 14.4316 | 0.010853 | 8.5331 |  | 0.0094431 | -5.5693 | 0.0099603 | -0.39696 | 0.0099908 | -0.092477 |  | Standard Error | 0.0026222
C3 | 0.01 | 0.0017 | 0.0022 | 0.011455 | 14.5529 | 0.011443 | 14.4316 | 0.010853 | 8.5331 |  | 0.0094431 | -5.5693 | 0.0099603 | -0.39696 | 0.0099908 | -0.092477 |  | Observations | 18
C3 | 0.01 | 0.0015 | 0.0023 | 0.01041 | 4.098 | 0.010403 | 4.0257 | 0.0097482 | -2.5176 |  | 0.0098515 | -1.4847 | 0.010342 | 3.4178 | 0.010428 | 4.282
C4 | 0.05 | 0.0083 | 0.0115 | 0.045957 | -8.0869 | 0.045816 | -8.3681 | 0.051542 | 3.0846 |  | 0.04743 | -5.14 | 0.045984 | -8.032 | 0.049541 | -0.91813 |  | ANOVA
C4 | 0.05 | 0.0079 | 0.0117 | 0.043866 | -12.2689 | 0.043731 | -12.5381 | 0.048843 | -2.3139 |  | 0.048247 | -3.5062 | 0.046771 | -6.4583 | 0.050366 | 0.73267 |  |  | df | SS | MS | F | Significance F
C4 | 0.05 | 0.008 | 0.0116 | 0.044388 | -11.2234 | 0.044252 | -11.4957 | 0.049515 | -0.96992 | Split | 0.047838 | -4.3231 | 0.046377 | -7.2453 | 0.049954 | -0.092465 | Split | Regression | 1 | 0.09079 | 0.09079 | 13203.9081 | 9.0462e-25
C5 | 0.1 | 0.0183 | 0.0253 | 0.098231 | -1.7689 | 0.098019 | -1.9812 | 0.094978 | -5.0219 |  | 0.1038 | 3.7977 | 0.10147 | 1.4727 | 0.098822 | -1.1775 |  | Residual | 16 | 0.00011002 | 6.876e-06
C5 | 0.1 | 0.0201 | 0.0257 | 0.10764 | 7.6406 | 0.10743 | 7.4311 | 0.10364 | 3.638 |  | 0.10543 | 5.4315 | 0.10312 | 3.1173 | 0.10038 | 0.37543 |  | Total | 17 | 0.0909
C5 | 0.1 | 0.0196 | 0.0258 | 0.10503 | 5.0268 | 0.10482 | 4.8161 | 0.10121 | 1.2092 |  | 0.10584 | 5.84 | 0.10353 | 3.5288 | 0.10076 | 0.76479
C6 | 0.2 | 0.0379 | 0.0481 | 0.20069 | 0.34468 | 0.20077 | 0.38291 | 0.20176 | 0.87976 |  | 0.19693 | -1.5365 | 0.19847 | -0.76284 | 0.19881 | -0.59553 |  |  | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
C6 | 0.2 | 0.0376 | 0.0483 | 0.19912 | -0.43944 | 0.19919 | -0.40556 | 0.19992 | -0.041045 |  | 0.19774 | -1.1281 | 0.19935 | -0.32272 | 0.19979 | -0.10535 |  | Intercept | 0.00045691 | 0.00081216 | 0.56259 | 0.58151 | -0.0012648 | 0.0021786 | -0.0012648 | 0.0021786 | NH3 Low curve coefficients:
C6 | 0.2 | 0.0373 | 0.0486 | 0.19755 | -1.2236 | 0.19761 | -1.194 | 0.19808 | -0.95863 |  | 0.19897 | -0.51538 | 0.20068 | 0.33795 | 0.20126 | 0.63162 |  | X Variable 1 | 4.0846 | 0.035547 | 114.9083 | 9.0462e-25 | 4.0093 | 4.16 | 4.0093 | 4.16 |  | intcpt | X coef. | X^2 coef.
[28 empty] | 0.0017 | 5.2244 | 94.06
PO4 Low curve coefficients:
[18 empty] | Record Int and Slope here: | [9 empty] | intcpt | X coef. | X^2 coef.
[18 empty] | NH3 Int. | 0.0025686 |  | PO4 Int. | 0.00045691 | [5 empty] | 0.0003 | 4.434 | -13.233
[18 empty] | NH3 Slope | 5.2275 |  | PO4 Slope | 4.0846
[27 empty] | Plot High Range calibrant data. |  |  | Enter coefficients below graph.
3. If errors exceed 10% from above, run quadratic fit.
(Use X=Peak Ht.; Y=Conc)
Enter equation parameters below graph.
NH3High curve coefficients:
[28 empty] | intcpt | X coef. | X^2 coef.
[28 empty] | 0.0201 | 3.4368 | 35.787
NH3 Quadratic curve coefficients
[21 empty] | Intercept | X coef. | X^2 coef. |  |  |  | PO4High curve coefficients:
[21 empty] | 0.0026 | 5.2006 | 0.7401 |  |  |  |  | intcpt | X coef. | X^2 coef.
[28 empty] | 0.0152 | 2.7371 | 22.456
PO4 Quadratic curve coefficients
[21 empty] | Intercept | X coef. | X^2 coef.
[21 empty] | 0.0016 | 3.7861 | 6.3809

## Apply Std Curve
|  |  | Sample Info |  |  |  | FASPac N-NH3 |  |  | FASPac P- PO4 |  |  | NH3 curve fit | PO4 curve fit | NH3 |  |  |  | PO4
SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | Peak# | Position | Identifier | Type | Raw Ht | Cor Ht | mg/L | Raw Ht | Cor Ht | mg/L | Results mg/L | Results mg/L | Curve-fitting parameters used: |  |  |  | Curve-fitting parameters used:
2019-06-10 | 1 | a | 26 | 1:25 | 1 | Unknown | 0.002 | 0.0011 | 0.008 | 0.0039 | 0.0042 | 0.0174 | 0.0075607 | 0.018689 | Copy from Std Curve Construction page. |  |  |  | Copy from Std Curve Construction page.
2019-06-10 | 1 | b | 27 | 1:26 | 2 | Unknown | 0.002 | 0.0014 | 0.0096 | 0.0039 | 0.0042 | 0.0176 | 0.0091985 | 0.018689
2019-06-10 | 1 | c | 28 | 1:27 | 3 | Unknown | 0.0015 | 0.001 | 0.0075 | 0.0044 | 0.0048 | 0.02 | 0.0070185 | 0.021278 | Use these for linear or quadratic fit. |  |  |  | Use these for linear or quadratic fit.
2019-06-10 | 2 | a | 29 | 1:28 | 4 | Unknown | 0.0023 | 0.0019 | 0.0122 | 0.0029 | 0.0033 | 0.0137 | 0.011966 | 0.014788 | If linear, enter 0 for X^2 coefficient. |  |  |  | If linear, enter 0 for X^2 coefficient.
2019-06-10 | 2 | b | 30 | 1:29 | 5 | Unknown | 0.0018 | 0.0014 | 0.0096 | 0.0034 | 0.0038 | 0.0158 | 0.0091985 | 0.016958
2019-06-10 | 2 | c | 31 | 1:30 | 6 | Unknown | 0.0012 | 0.0009 | 0.007 | 0.0047 | 0.0052 | 0.0215 | 0.0064781 | 0.022999 | Intercept: | X Coef. | X^2 Coef: |  | Intercept: | X Coef. | X^2 Coef:
2019-06-10 | 3 | a | 32 | 1:31 | 7 | Unknown | 0.0012 | 0.001 | 0.0075 | 0.0036 | 0.004 | 0.0168 | 0.0070185 | 0.017824 | -0.00555 | 6.4955 | 0 |  | -0.00182 | 3.6968 | 0
2019-06-10 | 3 | b | 33 | 1:32 | 8 | Unknown | 0.0015 | 0.0013 | 0.0091 | 0.0032 | 0.0037 | 0.0154 | 0.0086507 | 0.016525
2019-06-10 | 3 | c | 35 | 1:33 | 9 | Unknown | 0.0015 | 0.0013 | 0.0095 | 0.0039 | 0.0044 | 0.0184 | 0.0086507 | 0.019553
2019-06-10 | 4 | a | 36 | 1:34 | 10 | Unknown | 0.0015 | 0.0014 | 0.0096 | 0.0018 | 0.0022 | 0.0095 | 0.0091985 | 0.0099908 | Use these for split curve: |  |  |  | Use these for split curve:
2019-06-10 | 4 | b | 37 | 1:35 | 11 | Unknown | 0.0021 | 0.002 | 0.0128 | 0.0049 | 0.0053 | 0.0222 | 0.012525 | 0.023428 | Low Range: |  |  |  | Low Range:
2019-06-10 | 4 | c | 38 | 1:36 | 12 | Unknown | 0.0014 | 0.0013 | 0.009 | 0.0041 | 0.0045 | 0.0189 | 0.0086507 | 0.019985 | Intercept: | X Coef. | X^2 Coef: |  | Intercept: | X Coef. | X^2 Coef:
2019-06-10 | 5 | a | 39 | 1:37 | 13 | Unknown | 0.0014 | 0.0013 | 0.0093 | 0.0035 | 0.004 | 0.0165 | 0.0086507 | 0.017824 | 0.0017 | 5.2244 | 94.06 |  | 0.0003 | 4.434 | -13.233
2019-06-10 | 5 | b | 40 | 1:38 | 14 | Unknown | 0.0015 | 0.0013 | 0.0095 | 0.0025 | 0.003 | 0.0125 | 0.0086507 | 0.013483
2019-06-10 | 5 | c | 41 | 1:39 | 15 | Unknown | 0.0022 | 0.0021 | 0.0132 | 0.0041 | 0.0045 | 0.0189 | 0.013086 | 0.019985
2019-06-10 | 6 | a | 42 | 1:40 | 16 | Unknown | 0.0009 | 0.0008 | 0.0064 | 0.0021 | 0.0026 | 0.0109 | 0.0059397 | 0.011739 | High Range: |  |  |  | High Range:
2019-06-10 | 6 | b | 43 | 1:41 | 17 | Unknown | 0.0009 | 0.0008 | 0.0065 | 0.0036 | 0.0041 | 0.0172 | 0.0059397 | 0.018257 | Intercept: | X Coef. | X^2 Coef: |  | Intercept: | X Coef. | X^2 Coef:
2019-06-10 | 6 | c | 44 | 1:42 | 18 | Unknown | 0.0015 | 0.0014 | 0.01 | 0.0043 | 0.0047 | 0.0197 | 0.0091985 | 0.020847 | 0.0201 | 3.4368 | 35.787 |  | 0.0152 | 2.7371 | 22.456
2019-06-10 | 7 | a | 46 | 1:43 | 19 | Unknown | 0.002 | 0.0019 | 0.0124 | 0.0048 | 0.0052 | 0.0218 | 0.011966 | 0.022999
2019-06-10 | 7 | b | 47 | 1:44 | 20 | Unknown | 0.0021 | 0.0019 | 0.0124 | 0.0025 | 0.003 | 0.0126 | 0.011966 | 0.013483 | Procedure for split curve: |  |  |  | Procedure for split curve:
2019-06-10 | 7 | c | 48 | 1:45 | 21 | Unknown | 0.0012 | 0.001 | 0.0076 | 0.0037 | 0.0041 | 0.0173 | 0.0070185 | 0.018257 | 1. Sort data on Row I: FASpac results, low to high. |  |  |  | 1. Sort data on Row I: FASpac results, low to high.
2019-06-10 | 8 | a | 49 | 1:46 | 22 | Unknown | 0.0009 | 0.0007 | 0.0062 | 0.0047 | 0.0052 | 0.0216 | 0.0054032 | 0.022999 | 2. Apply low range equation. |  |  |  | 2. Apply low range equation.
2019-06-10 | 8 | b | 50 | 1:47 | 23 | Unknown | 0.0017 | 0.0015 | 0.0104 | 0.0049 | 0.0054 | 0.0223 | 0.0097482 | 0.023858 | 3. Apply high range equation. |  |  |  | 3. Apply high range equation.
2019-06-10 | 8 | c | 51 | 1:48 | 24 | Unknown | 0.0016 | 0.0014 | 0.0096 | 0.0044 | 0.0048 | 0.0201 | 0.0091985 | 0.021278 | 4. Note here peak height where split occurs: |  |  | 0.05 | 4. Note here peak height where split occurs: |  |  | 0.05
2019-06-10 | 9 | a | 52 | 1:49 | 25 | Unknown | 0.0011 | 0.0008 | 0.0067 | 0.0024 | 0.0029 | 0.0121 | 0.0059397 | 0.013047 | 4. Sort back based on Row D: Peak Number |  |  |  | 4. Sort back based on Row D: Peak Number
2019-06-10 | 9 | b | 53 | 1:50 | 26 | Unknown | 0.0014 | 0.0012 | 0.0086 | 0.0042 | 0.0047 | 0.0197 | 0.0081047 | 0.020847
2019-06-10 | 9 | c | 54 | 1:51 | 27 | Unknown | 0.0009 | 0.0006 | 0.0054 | 0.0038 | 0.0043 | 0.0179 | 0.0048685 | 0.019122
2019-06-10 | 10 | a | 55 | 1:52 | 28 | Unknown | 0.0005 | 0.0002 | 0.0035 | 0.0034 | 0.0039 | 0.0163 | 0.0027486 | 0.017391
2019-06-10 | 10 | b | 57 | 1:53 | 29 | Unknown | 0.001 | 0.0008 | 0.0064 | 0.0055 | 0.006 | 0.0248 | 0.0059397 | 0.026428
2019-06-10 | 10 | c | 58 | 1:54 | 30 | Unknown | 0.0007 | 0.0004 | 0.0046 | 0.0042 | 0.0047 | 0.0195 | 0.0038048 | 0.020847
2019-06-10 | 11 | a | 59 | 1:55 | 31 | Unknown | 0.0012 | 0.001 | 0.0077 | 0.0032 | 0.0037 | 0.0157 | 0.0070185 | 0.016525
2019-06-10 | 11 | b | 60 | 1:56 | 32 | Unknown | 0.0012 | 0.001 | 0.0075 | 0.0026 | 0.0032 | 0.0133 | 0.0070185 | 0.014353
2019-06-10 | 11 | c | 61 | 1:57 | 33 | Unknown | 0.0007 | 0.0005 | 0.0049 | 0.0022 | 0.0028 | 0.0117 | 0.0043357 | 0.012611
2019-06-10 | 12 | a | 62 | 1:58 | 34 | Unknown | 0.0013 | 0.0011 | 0.0083 | 0.0018 | 0.0024 | 0.0102 | 0.0075607 | 0.010865
2019-06-10 | 12 | b | 63 | 1:59 | 35 | Unknown | 0.0007 | 0.0006 | 0.0055 | 0.0034 | 0.004 | 0.0165 | 0.0048685 | 0.017824
2019-06-10 | 12 | c | 64 | 1:60 | 36 | Unknown | 0.0011 | 0.001 | 0.0077 | 0.0023 | 0.0029 | 0.0121 | 0.0070185 | 0.013047
2019-06-10 | 13 | a | 65 | 2:1 | 37 | Unknown | 0.0012 | 0.0011 | 0.0084 | 0.0009 | 0.0015 | 0.0064 | 0.0075607 | 0.0069212
2019-06-10 | 13 | b | 66 | 2:2 | 38 | Unknown | 0.0007 | 0.0006 | 0.0057 | 0.0017 | 0.0023 | 0.0098 | 0.0048685 | 0.010428
2019-06-10 | 13 | c | 68 | 2:3 | 39 | Unknown | 0.0008 | 0.0008 | 0.0065 | 0.0011 | 0.0018 | 0.0076 | 0.0059397 | 0.0082383
2019-06-25 | 1 | a | 69 | 2:4 | 40 | Unknown | 0.0012 | 0.0012 | 0.0087 | 0.0042 | 0.0048 | 0.02 | 0.0081047 | 0.021278
2019-06-25 | 1 | b | 70 | 2:5 | 41 | Unknown | 0.0008 | 0.0009 | 0.007 | 0.0039 | 0.0045 | 0.0186 | 0.0064781 | 0.019985
2019-06-25 | 1 | c | 71 | 2:6 | 42 | Unknown | 0.001 | 0.001 | 0.0078 | 0.0044 | 0.005 | 0.021 | 0.0070185 | 0.022139
2019-06-25 | 2 | a | 72 | 2:7 | 43 | Unknown | 0.0014 | 0.0015 | 0.01 | 0.0027 | 0.0033 | 0.014 | 0.0097482 | 0.014788
2019-06-25 | 2 | b | 73 | 2:8 | 44 | Unknown | 0.8956 | 0.8956 | 4.6968 | 1.3846 | 1.3852 | 5.6692 | 80.1261 | -19.2489
2019-06-25 | 2 | c | 74 | 2:9 | 45 | Unknown | 0.0013 | -0.0105 | -0.0527 | 0.0037 | 0 | 0.0004 | -0.042786 | 0.0003
2019-06-25 | 3 | a | 75 | 2:10 | 46 | Unknown | 0.0008 | 0.001 | 0.0076 | 0.0045 | 0.0051 | 0.0211 | 0.0070185 | 0.022569
2019-06-25 | 3 | b | 76 | 2:11 | 47 | Unknown | 0.001 | 0.001 | 0.0079 | 0.0041 | 0.0047 | 0.0196 | 0.0070185 | 0.020847
2019-06-25 | 3 | c | 77 | 2:12 | 48 | Unknown | 0.0015 | 0.0016 | 0.0108 | 0.0041 | 0.0047 | 0.0194 | 0.0103 | 0.020847
2019-06-25 | 4 | a | 79 | 2:13 | 49 | Unknown | 0.0017 | 0.0017 | 0.0115 | 0.0034 | 0.004 | 0.0167 | 0.010853 | 0.017824
2019-06-25 | 4 | b | 80 | 2:14 | 50 | Unknown | 0.0004 | 0.0004 | 0.0044 | -0.0006 | -0.0001 | 0.0001 | 0.0038048 | -0.00014353
2019-06-25 | 4 | c | 83 | 2:15 | 51 | Unknown | 0.0033 | 0.003 | 0.0183 | 0.0008 | 0.0014 | 0.0059 | 0.01822 | 0.0064817
2019-06-25 | 5 | a | 84 | 2:16 | 52 | Unknown | 0.001 | 0.0009 | 0.0071 | 0.0022 | 0.0028 | 0.0119 | 0.0064781 | 0.012611
2019-06-25 | 5 | b | 85 | 2:17 | 53 | Unknown | 0.0012 | 0.0011 | 0.008 | 0.0026 | 0.0033 | 0.0138 | 0.0075607 | 0.014788
2019-06-25 | 5 | c | 86 | 2:18 | 54 | Unknown | 0.0009 | 0.0008 | 0.0065 | 0.0025 | 0.0031 | 0.0131 | 0.0059397 | 0.013918
2019-06-25 | 6 | a | 87 | 2:19 | 55 | Unknown | 0.0006 | 0.0005 | 0.005 | 0.0013 | 0.002 | 0.0085 | 0.0043357 | 0.0091151
2019-06-25 | 6 | b | 88 | 2:20 | 56 | Unknown | 0.0006 | 0.0005 | 0.0051 | 0.0009 | 0.0016 | 0.0068 | 0.0043357 | 0.0073605
2019-06-25 | 6 | c | 90 | 2:21 | 57 | Unknown | 0.0013 | 0.0012 | 0.0085 | 0.0006 | 0.0013 | 0.0056 | 0.0081047 | 0.0060418
2019-06-25 | 7 | a | 91 | 2:22 | 58 | Unknown | 0.0009 | 0.0008 | 0.0064 | 0.0009 | 0.0016 | 0.007 | 0.0059397 | 0.0073605
2019-06-25 | 7 | b | 92 | 2:23 | 59 | Unknown | 0.9384 | 0.9382 | 4.9202 | 1.6165 | 1.6172 | 6.6186 | 87.6967 | -27.4378
2019-06-25 | 7 | c | 93 | 2:24 | 60 | Unknown | 0.0017 | -0.0108 | -0.0544 | 0.0009 | -0.0034 | -0.0135 | -0.043752 | -0.014929
2019-06-25 | 8 | a | 94 | 2:25 | 61 | Unknown | 0.0013 | 0.0013 | 0.0092 | 0.0014 | 0.0021 | 0.0091 | 0.0086507 | 0.009553
2019-06-25 | 8 | b | 95 | 2:26 | 62 | Unknown | 0.9161 | 0.9159 | 4.8031 | 1.5517 | 1.5523 | 6.353 | 83.6911 | -25.0035
2019-06-25 | 8 | c | 96 | 2:27 | 63 | Unknown | 0.0005 | -0.0117 | -0.0591 | -0.0007 | -0.0048 | -0.0194 | -0.04655 | -0.021288
2019-06-25 | 9 | a | 97 | 2:28 | 64 | Unknown | 0.0005 | 0.0005 | 0.005 | 0.0019 | 0.0026 | 0.011 | 0.0043357 | 0.011739
2019-06-25 | 9 | b | 98 | 2:29 | 65 | Unknown | 0.0005 | 0.0004 | 0.0044 | -0.0005 | 0.0001 | 0.0008 | 0.0038048 | 0.00074327
2019-06-25 | 9 | c | 99 | 2:30 | 66 | Unknown | 0.0006 | 0.0005 | 0.0049 | 0.0023 | 0.003 | 0.0125 | 0.0043357 | 0.013483
2019-06-25 | 10 | a | 101 | 2:31 | 67 | Unknown | 0.0003 | 0.0001 | 0.0031 | 0.004 | 0.0047 | 0.0194 | 0.0022234 | 0.020847
2019-06-25 | 10 | b | 102 | 2:32 | 68 | Unknown | 0.0002 | 0.0001 | 0.0027 | 0.0037 | 0.0044 | 0.0182 | 0.0022234 | 0.019553
2019-06-25 | 10 | c | 103 | 2:33 | 69 | Unknown | 0.0007 | 0.0005 | 0.0052 | 0.0031 | 0.0037 | 0.0157 | 0.0043357 | 0.016525
2019-06-25 | 11 | a | 104 | 2:34 | 70 | Unknown | 0.0006 | 0.0004 | 0.0047 | 0.0026 | 0.0032 | 0.0135 | 0.0038048 | 0.014353
2019-06-25 | 11 | b | 105 | 2:35 | 71 | Unknown | 0.0078 | 0.0076 | 0.0423 | 0.0018 | 0.0024 | 0.0103 | 0.046838 | 0.010865
2019-06-25 | 11 | c | 106 | 2:36 | 72 | Unknown | 0.0008 | 0.0006 | 0.0056 | 0.0019 | 0.0024 | 0.0103 | 0.0048685 | 0.010865
2019-06-25 | 12 | a | 107 | 2:37 | 73 | Unknown | 0.0019 | 0.0017 | 0.0115 | 0.0016 | 0.0021 | 0.0091 | 0.010853 | 0.009553
2019-06-25 | 12 | b | 108 | 2:38 | 74 | Unknown | 0.9226 | 0.9225 | 4.8376 | 1.5634 | 1.564 | 6.4006 | 84.8669 | -25.4341
2019-06-25 | 12 | c | 109 | 2:39 | 75 | Unknown | 0.0012 | -0.0111 | -0.0559 | 0.0018 | -0.0024 | -0.0096 | -0.044702 | -0.010418
2019-06-25 | 13 | a | 110 | 2:40 | 76 | Unknown | 0.0004 | 0.0004 | 0.0045 | 0.0007 | 0.0013 | 0.0055 | 0.0038048 | 0.0060418
2019-06-25 | 13 | b | 112 | 2:41 | 77 | Unknown | 0.0027 | 0.0025 | 0.0156 | -0.0006 | 0 | 0.0003 | 0.015349 | 0.0003
2019-06-25 | 13 | c | 113 | 2:42 | 78 | Unknown | 0.0008 | 0.0006 | 0.0054 | 0.0006 | 0.0012 | 0.0052 | 0.0048685 | 0.0056017
2019-07-09 | 1 | a | 114 | 2:43 | 79 | Unknown | 0.002 | 0.0018 | 0.012 | 0.0047 | 0.0052 | 0.0217 | 0.011409 | 0.022999
2019-07-09 | 1 | b | 115 | 2:44 | 80 | Unknown | 0.0013 | 0.001 | 0.0079 | 0.0049 | 0.0054 | 0.0225 | 0.0070185 | 0.023858
2019-07-09 | 1 | c | 116 | 2:45 | 81 | Unknown | 0.0037 | 0.0035 | 0.0206 | 0.0057 | 0.0063 | 0.0259 | 0.021138 | 0.027709
2019-07-09 | 2 | a | 119 | 2:46 | 82 | Unknown | 0.0019 | 0.0013 | 0.0094 | 0.0061 | 0.0066 | 0.0273 | 0.0086507 | 0.028988
2019-07-09 | 2 | b | 120 | 2:47 | 83 | Unknown | 0.0016 | 0.0012 | 0.0088 | 0.0056 | 0.0062 | 0.0258 | 0.0081047 | 0.027282
2019-07-09 | 2 | c | 121 | 2:48 | 84 | Unknown | 0.0018 | 0.0014 | 0.01 | 0.0049 | 0.0054 | 0.0226 | 0.0091985 | 0.023858
2019-07-09 | 3 | a | 123 | 2:49 | 85 | Unknown | 0.0015 | 0.0012 | 0.0085 | 0.0051 | 0.0057 | 0.0237 | 0.0081047 | 0.025144
2019-07-09 | 3 | b | 124 | 2:50 | 86 | Unknown | 0.0014 | 0.001 | 0.0079 | 0.0042 | 0.0049 | 0.0203 | 0.0070185 | 0.021709
2019-07-09 | 3 | c | 125 | 2:51 | 87 | Unknown | 0.0016 | 0.0013 | 0.0092 | 0.0039 | 0.0046 | 0.019 | 0.0086507 | 0.020416
2019-07-09 | 4 | a | 126 | 2:52 | 88 | Unknown | 0.0016 | 0.0012 | 0.0089 | 0.0039 | 0.0046 | 0.0191 | 0.0081047 | 0.020416
2019-07-09 | 4 | b | 127 | 2:53 | 89 | Unknown | 0.0011 | 0.0008 | 0.0066 | 0.003 | 0.0037 | 0.0154 | 0.0059397 | 0.016525
2019-07-09 | 4 | c | 128 | 2:54 | 90 | Unknown | 0.0011 | 0.0009 | 0.0069 | 0.0037 | 0.0043 | 0.0181 | 0.0064781 | 0.019122
2019-07-09 | 5 | a | 129 | 2:55 | 91 | Unknown | 0.0013 | 0.001 | 0.0079 | 0.003 | 0.0037 | 0.0154 | 0.0070185 | 0.016525
2019-07-09 | 5 | b | 130 | 2:56 | 92 | Unknown | 0.0012 | 0.0009 | 0.0073 | 0.0028 | 0.0035 | 0.0147 | 0.0064781 | 0.015657
2019-07-09 | 5 | c | 131 | 2:57 | 93 | Unknown | 0.0015 | 0.0013 | 0.0091 | 0.0022 | 0.0029 | 0.012 | 0.0086507 | 0.013047
2019-07-09 | 6 | a | 132 | 2:58 | 94 | Unknown | 0.0011 | 0.0008 | 0.0068 | 0.0019 | 0.0026 | 0.0111 | 0.0059397 | 0.011739
2019-07-09 | 6 | b | 134 | 2:59 | 95 | Unknown | 0.0011 | 0.0008 | 0.0068 | 0.0016 | 0.0023 | 0.0096 | 0.0059397 | 0.010428
2019-07-09 | 6 | c | 135 | 2:60 | 96 | Unknown | 0.0016 | 0.0014 | 0.0098 | 0.0014 | 0.002 | 0.0087 | 0.0091985 | 0.0091151
2019-07-09 | 7 | a | 136 | 3:1 | 97 | Unknown | 0.0009 | 0.0006 | 0.0056 | 0.0015 | 0.0022 | 0.0094 | 0.0048685 | 0.0099908
2019-07-09 | 7 | b | 137 | 3:2 | 98 | Unknown | 0.0008 | 0.0006 | 0.0054 | 0.0022 | 0.0029 | 0.0122 | 0.0048685 | 0.013047
2019-07-09 | 7 | c | 138 | 3:3 | 99 | Unknown | 0.0024 | 0.0021 | 0.0135 | 0.0018 | 0.0024 | 0.0102 | 0.013086 | 0.010865
2019-07-09 | 8 | a | 139 | 3:4 | 100 | Unknown | 0.0007 | 0.0004 | 0.0044 | 0.0035 | 0.0042 | 0.0175 | 0.0038048 | 0.018689
2019-07-09 | 8 | b | 140 | 3:5 | 101 | Unknown | 0.0007 | 0.0004 | 0.0045 | 0.0029 | 0.0035 | 0.0146 | 0.0038048 | 0.015657
2019-07-09 | 8 | c | 141 | 3:6 | 102 | Unknown | 0.0008 | 0.0005 | 0.005 | 0.0036 | 0.0042 | 0.0175 | 0.0043357 | 0.018689
2019-07-09 | 9 | a | 142 | 3:7 | 103 | Unknown | 0.0009 | 0.0006 | 0.0056 | 0.0031 | 0.0037 | 0.0155 | 0.0048685 | 0.016525
2019-07-09 | 9 | b | 143 | 3:8 | 104 | Unknown | 0.0009 | 0.0005 | 0.0053 | 0.0034 | 0.004 | 0.0166 | 0.0043357 | 0.017824
2019-07-09 | 9 | c | 145 | 3:9 | 105 | Unknown | 0.001 | 0.0007 | 0.0061 | 0.0025 | 0.0031 | 0.0129 | 0.0054032 | 0.013918
2019-07-09 | 10 | a | 146 | 3:10 | 106 | Unknown | 0.0029 | 0.0025 | 0.0156 | 0.0041 | 0.0047 | 0.0197 | 0.015349 | 0.020847
2019-07-09 | 10 | b | 147 | 3:11 | 107 | Unknown | 0.0012 | 0.0009 | 0.0071 | 0.0036 | 0.0042 | 0.0176 | 0.0064781 | 0.018689
2019-07-09 | 10 | c | 148 | 3:12 | 108 | Unknown | 0.001 | 0.0007 | 0.0062 | 0.0038 | 0.0044 | 0.0184 | 0.0054032 | 0.019553
2019-07-09 | 11 | a | 149 | 3:13 | 109 | Unknown | 0.0009 | 0.0006 | 0.0056 | 0.003 | 0.0036 | 0.0152 | 0.0048685 | 0.016091
2019-07-09 | 11 | b | 150 | 3:14 | 110 | Unknown | 0.0011 | 0.0008 | 0.0066 | 0.0026 | 0.0032 | 0.0134 | 0.0059397 | 0.014353
2019-07-09 | 11 | c | 151 | 3:15 | 111 | Unknown | 0.0011 | 0.0007 | 0.0063 | 0.003 | 0.0036 | 0.015 | 0.0054032 | 0.016091
2019-07-09 | 12 | a | 152 | 3:16 | 112 | Unknown | 0.0013 | 0.001 | 0.0078 | 0.0013 | 0.0019 | 0.0082 | 0.0070185 | 0.0086768
2019-07-09 | 12 | b | 153 | 3:17 | 113 | Unknown | 0.0012 | 0.0009 | 0.0071 | 0.001 | 0.0016 | 0.007 | 0.0064781 | 0.0073605
2019-07-09 | 12 | c | 154 | 3:18 | 114 | Unknown | 0.0011 | 0.0008 | 0.0067 | 0.0019 | 0.0025 | 0.0106 | 0.0059397 | 0.011302
2019-07-09 | 13 | a | 156 | 3:19 | 115 | Unknown | 0.0016 | 0.0013 | 0.0094 | 0.0004 | 0.001 | 0.0043 | 0.0086507 | 0.0047208
2019-07-09 | 13 | b | 157 | 3:20 | 116 | Unknown | 0.0011 | 0.0008 | 0.0067 | 0.0002 | 0.0007 | 0.0034 | 0.0059397 | 0.0033973
2019-07-09 | 13 | c | 158 | 3:21 | 117 | Unknown | 0.0019 | 0.0017 | 0.0112 | 0.0003 | 0.0008 | 0.0037 | 0.010853 | 0.0038387

## Final Results
SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | N-NH3 mg/L | P-PO4 mg/L
2019-06-10 | 1 | a | 0.0075607 | 0.018689
2019-06-10 | 1 | b | 0.0091985 | 0.018689
2019-06-10 | 1 | c | 0.0070185 | 0.021278
2019-06-10 | 2 | a | 0.011966 | 0.014788
2019-06-10 | 2 | b | 0.0091985 | 0.016958
2019-06-10 | 2 | c | 0.0064781 | 0.022999
2019-06-10 | 3 | a | 0.0070185 | 0.017824
2019-06-10 | 3 | b | 0.0086507 | 0.016525
2019-06-10 | 3 | c | 0.0086507 | 0.019553
2019-06-10 | 4 | a | 0.0091985 | 0.0099908
2019-06-10 | 4 | b | 0.012525 | 0.023428
2019-06-10 | 4 | c | 0.0086507 | 0.019985
2019-06-10 | 5 | a | 0.0086507 | 0.017824
2019-06-10 | 5 | b | 0.0086507 | 0.013483
2019-06-10 | 5 | c | 0.013086 | 0.019985
2019-06-10 | 6 | a | 0.0059397 | 0.011739
2019-06-10 | 6 | b | 0.0059397 | 0.018257
2019-06-10 | 6 | c | 0.0091985 | 0.020847
2019-06-10 | 7 | a | 0.011966 | 0.022999
2019-06-10 | 7 | b | 0.011966 | 0.013483
2019-06-10 | 7 | c | 0.0070185 | 0.018257
2019-06-10 | 8 | a | 0.0054032 | 0.022999
2019-06-10 | 8 | b | 0.0097482 | 0.023858
2019-06-10 | 8 | c | 0.0091985 | 0.021278
2019-06-10 | 9 | a | 0.0059397 | 0.013047
2019-06-10 | 9 | b | 0.0081047 | 0.020847
2019-06-10 | 9 | c | 0.0048685 | 0.019122
2019-06-10 | 10 | a | 0.0027486 | 0.017391
2019-06-10 | 10 | b | 0.0059397 | 0.026428
2019-06-10 | 10 | c | 0.0038048 | 0.020847
2019-06-10 | 11 | a | 0.0070185 | 0.016525
2019-06-10 | 11 | b | 0.0070185 | 0.014353
2019-06-10 | 11 | c | 0.0043357 | 0.012611
2019-06-10 | 12 | a | 0.0075607 | 0.010865
2019-06-10 | 12 | b | 0.0048685 | 0.017824
2019-06-10 | 12 | c | 0.0070185 | 0.013047
2019-06-10 | 13 | a | 0.0075607 | 0.0069212
2019-06-10 | 13 | b | 0.0048685 | 0.010428
2019-06-10 | 13 | c | 0.0059397 | 0.0082383
2019-06-25 | 1 | a | 0.0081047 | 0.021278
2019-06-25 | 1 | b | 0.0064781 | 0.019985
2019-06-25 | 1 | c | 0.0070185 | 0.022139
2019-06-25 | 2 | a | 0.0097482 | 0.014788
2019-06-25 | 2 | b | 80.1261 | -19.2489 | Outlier
2019-06-25 | 2 | c | -0.042786 | 0.0003
2019-06-25 | 3 | a | 0.0070185 | 0.022569
2019-06-25 | 3 | b | 0.0070185 | 0.020847
2019-06-25 | 3 | c | 0.0103 | 0.020847
2019-06-25 | 4 | a | 0.010853 | 0.017824
2019-06-25 | 4 | b | 0.0038048 | -0.00014353
2019-06-25 | 4 | c | 0.01822 | 0.0064817
2019-06-25 | 5 | a | 0.0064781 | 0.012611
2019-06-25 | 5 | b | 0.0075607 | 0.014788
2019-06-25 | 5 | c | 0.0059397 | 0.013918
2019-06-25 | 6 | a | 0.0043357 | 0.0091151
2019-06-25 | 6 | b | 0.0043357 | 0.0073605
2019-06-25 | 6 | c | 0.0081047 | 0.0060418
2019-06-25 | 7 | a | 0.0059397 | 0.0073605
2019-06-25 | 7 | b | 87.6967 | -27.4378 | Outlier
2019-06-25 | 7 | c | -0.043752 | -0.014929
2019-06-25 | 8 | a | 0.0086507 | 0.009553
2019-06-25 | 8 | b | 83.6911 | -25.0035 | Outlier
2019-06-25 | 8 | c | -0.04655 | -0.021288
2019-06-25 | 9 | a | 0.0043357 | 0.011739
2019-06-25 | 9 | b | 0.0038048 | 0.00074327
2019-06-25 | 9 | c | 0.0043357 | 0.013483
2019-06-25 | 10 | a | 0.0022234 | 0.020847
2019-06-25 | 10 | b | 0.0022234 | 0.019553
2019-06-25 | 10 | c | 0.0043357 | 0.016525
2019-06-25 | 11 | a | 0.0038048 | 0.014353
2019-06-25 | 11 | b | 0.046838 | 0.010865
2019-06-25 | 11 | c | 0.0048685 | 0.010865
2019-06-25 | 12 | a | 0.010853 | 0.009553
2019-06-25 | 12 | b | 84.8669 | -25.4341 | Outlier
2019-06-25 | 12 | c | -0.044702 | -0.010418
2019-06-25 | 13 | a | 0.0038048 | 0.0060418
2019-06-25 | 13 | b | 0.015349 | 0.0003
2019-06-25 | 13 | c | 0.0048685 | 0.0056017
2019-07-09 | 1 | a | 0.011409 | 0.022999
2019-07-09 | 1 | b | 0.0070185 | 0.023858
2019-07-09 | 1 | c | 0.021138 | 0.027709
2019-07-09 | 2 | a | 0.0086507 | 0.028988
2019-07-09 | 2 | b | 0.0081047 | 0.027282
2019-07-09 | 2 | c | 0.0091985 | 0.023858
2019-07-09 | 3 | a | 0.0081047 | 0.025144
2019-07-09 | 3 | b | 0.0070185 | 0.021709
2019-07-09 | 3 | c | 0.0086507 | 0.020416
2019-07-09 | 4 | a | 0.0081047 | 0.020416
2019-07-09 | 4 | b | 0.0059397 | 0.016525
2019-07-09 | 4 | c | 0.0064781 | 0.019122
2019-07-09 | 5 | a | 0.0070185 | 0.016525
2019-07-09 | 5 | b | 0.0064781 | 0.015657
2019-07-09 | 5 | c | 0.0086507 | 0.013047
2019-07-09 | 6 | a | 0.0059397 | 0.011739
2019-07-09 | 6 | b | 0.0059397 | 0.010428
2019-07-09 | 6 | c | 0.0091985 | 0.0091151
2019-07-09 | 7 | a | 0.0048685 | 0.0099908
2019-07-09 | 7 | b | 0.0048685 | 0.013047
2019-07-09 | 7 | c | 0.013086 | 0.010865
2019-07-09 | 8 | a | 0.0038048 | 0.018689
2019-07-09 | 8 | b | 0.0038048 | 0.015657
2019-07-09 | 8 | c | 0.0043357 | 0.018689
2019-07-09 | 9 | a | 0.0048685 | 0.016525
2019-07-09 | 9 | b | 0.0043357 | 0.017824
2019-07-09 | 9 | c | 0.0054032 | 0.013918
2019-07-09 | 10 | a | 0.015349 | 0.020847
2019-07-09 | 10 | b | 0.0064781 | 0.018689
2019-07-09 | 10 | c | 0.0054032 | 0.019553
2019-07-09 | 11 | a | 0.0048685 | 0.016091
2019-07-09 | 11 | b | 0.0059397 | 0.014353
2019-07-09 | 11 | c | 0.0054032 | 0.016091
2019-07-09 | 12 | a | 0.0070185 | 0.0086768
2019-07-09 | 12 | b | 0.0064781 | 0.0073605
2019-07-09 | 12 | c | 0.0059397 | 0.011302
2019-07-09 | 13 | a | 0.0086507 | 0.0047208
2019-07-09 | 13 | b | 0.0059397 | 0.0033973
2019-07-09 | 13 | c | 0.010853 | 0.0038387