## Results and Sample ID
[5 empty] | Run Name: | LTREB NO3_AP2_2019_10_25
[5 empty] | Run Date: | 2019-10-25
[5 empty] | Operator: | RE
|  |  |  | Sample Info |  |  |  | FASPac NO3-NO2 |  |  | Calibrant | Conc., mg/L
Row | SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | Peak# | Position | Identifier | Type | Raw Ht | Cor Ht | mg/L | C1 | 0.001
|  |  |  | 1 | 1:1 | SYNC | SYNC | 0.3927 | 0.3704 | 0.478 | C2 | 0.005
|  |  |  | 2 | 1:2 | CO | Carry over | 0.0246 | 0.0023 | 0.003 | C3 | 0.01
|  |  |  | 3 | 1:3 | W | Wash | 0.0224 | 0 | 0 | C4 | 0.05
|  |  |  | 4 | 1:4 | NO3 | Unknown | 0.1024 | 0.0803 | 0.104 | C5 | 0.1
|  |  |  | 5 | 1:5 | NOX% | NO3 check | 0.1032 | 0.081 | 0.104 | C6 | 0.5
[6 empty] | NO3 Efficiency |  |  |  | 99.22%
|  |  |  | 6 | 1:3 | W | Wash | 0.0212 | 0 | 0
|  |  |  | 7 | 1:7 | C1 | Calibrant | 0.022 | 0.0006 | 0.001
|  |  |  | 8 | 1:7 | C1 | Calibrant | 0.0219 | 0.0002 | 0
|  |  |  | 9 | 1:7 | C1 | Calibrant | 0.0219 | -0.0002 | 0
|  |  |  | 10 | 1:8 | C2 | Calibrant | 0.0266 | 0.0042 | 0.005
|  |  |  | 11 | 1:8 | C2 | Calibrant | 0.0266 | 0.0038 | 0.005
|  |  |  | 12 | 1:8 | C2 | Calibrant | 0.0266 | 0.0036 | 0.005
|  |  |  | 13 | 1:9 | C3 | Calibrant | 0.0305 | 0.0071 | 0.009
|  |  |  | 14 | 1:9 | C3 | Calibrant | 0.0318 | 0.0081 | 0.01
|  |  |  | 15 | 1:9 | C3 | Calibrant | 0.0318 | 0.0077 | 0.01
|  |  |  | 16 | 1:10 | C4 | Calibrant | 0.0641 | 0.0397 | 0.051
|  |  |  | 17 | 1:10 | C4 | Calibrant | 0.0641 | 0.0391 | 0.05
|  |  |  | 18 | 1:10 | C4 | Calibrant | 0.0646 | 0.0393 | 0.051
|  |  |  | 19 | 1:11 | C5 | Calibrant | 0.1032 | 0.0775 | 0.1
|  |  |  | 20 | 1:11 | C5 | Calibrant | 0.1038 | 0.0776 | 0.1
|  |  |  | 21 | 1:11 | C5 | Calibrant | 0.1046 | 0.0781 | 0.101
|  |  |  | 22 | 1:12 | C6 | Calibrant | 0.4161 | 0.3892 | 0.502
|  |  |  | 23 | 1:12 | C6 | Calibrant | 0.416 | 0.3869 | 0.499
|  |  |  | 24 | 1:12 | C6 | Calibrant | 0.4162 | 0.3868 | 0.499
|  |  |  | 25 | 1:3 | W | Wash | 0.0274 | 0 | 0
|  |  |  | 26 | 1:4 | NO3 | Unknown | 0.1057 | 0.0785 | 0.101
|  |  |  | 27 | 1:5 | NOX% | NO3 check | 0.106 | 0.0785 | 0.101
[6 empty] | NO3 Efficiency |  |  |  | 100.00%
|  |  |  | 28 | 1:13 | CC1 | QC 2 | 0.0419 | 0.0146 | 0.019
|  |  |  | 29 | 1:14 | CC2 | Unknown | 0.1845 | 0.1578 | 0.204
|  |  |  | 30 | 1:3 | W | Wash | 0.0263 | 0 | 0
| 2019-07-23 | 1 | a | 31 | 1:25 | 1 | Unknown | 0.0485 | 0.0223 | 0.029
| 2019-07-23 | 1 | b | 32 | 1:26 | 2 | Unknown | 0.0449 | 0.0186 | 0.024
| 2019-07-23 | 1 | c | 33 | 1:27 | 3 | Unknown | 0.0492 | 0.023 | 0.03
| 2019-07-23 | 2 | a | 34 | 1:28 | 4 | Unknown | 0.0444 | 0.0183 | 0.024
| 2019-07-23 | 2 | b | 35 | 1:29 | 5 | Unknown | 0.0432 | 0.0172 | 0.022
| 2019-07-23 | 2 | c | 37 | 1:30 | 6 | Unknown | 0.0439 | 0.0181 | 0.023
| 2019-07-23 | 3 | a | 38 | 1:31 | 7 | Unknown | 0.0415 | 0.0156 | 0.02
| 2019-07-23 | 3 | b | 39 | 1:32 | 8 | Unknown | 0.043 | 0.017 | 0.022
| 2019-07-23 | 3 | c | 40 | 1:33 | 9 | Unknown | 0.0431 | 0.0171 | 0.022
| 2019-07-23 | 4 | a | 41 | 1:34 | 10 | Unknown | 0.0353 | 0.0093 | 0.012
| 2019-07-23 | 4 | b | 42 | 1:35 | 11 | Unknown | 0.0354 | 0.0095 | 0.012
| 2019-07-23 | 4 | c | 43 | 1:36 | 12 | Unknown | 0.0345 | 0.0085 | 0.011
| 2019-07-23 | 5 | a | 44 | 1:37 | 13 | Unknown | 0.0352 | 0.0092 | 0.012
| 2019-07-23 | 5 | b | 45 | 1:38 | 14 | Unknown | 0.0352 | 0.0092 | 0.012
| 2019-07-23 | 5 | c | 46 | 1:39 | 15 | Unknown | 0.0352 | 0.0092 | 0.012
| 2019-07-23 | 6 | a | 50 | 1:40 | 16 | Unknown | 0.0445 | 0.0179 | 0.023
| 2019-07-23 | 6 | b | 51 | 1:41 | 17 | Unknown | 0.0431 | 0.017 | 0.022
| 2019-07-23 | 6 | c | 52 | 1:42 | 18 | Unknown | 0.0435 | 0.0173 | 0.022
| 2019-07-23 | 7 | a | 53 | 1:43 | 19 | Unknown | 0.0288 | 0.0026 | 0.003
| 2019-07-23 | 7 | b | 54 | 1:44 | 20 | Unknown | 0.0284 | 0.0022 | 0.003
| 2019-07-23 | 7 | c | 55 | 1:45 | 21 | Unknown | 0.0292 | 0.0031 | 0.004
| 2019-07-23 | 8 | a | 56 | 1:46 | 22 | Unknown | 0.029 | 0.0029 | 0.004
| 2019-07-23 | 8 | b | 57 | 1:47 | 23 | Unknown | 0.0284 | 0.0023 | 0.003
| 2019-07-23 | 8 | c | 59 | 1:48 | 24 | Unknown | 0.0281 | 0.0019 | 0.003
| 2019-07-23 | 9 | a | 60 | 1:49 | 25 | Unknown | 0.0278 | 0.0016 | 0.002
| 2019-07-23 | 9 | b | 61 | 1:50 | 26 | Unknown | 0.0277 | 0.0014 | 0.002
| 2019-07-23 | 9 | c | 62 | 1:51 | 27 | Unknown | 0.028 | 0.0017 | 0.002
| 2019-07-23 | 10 | a | 63 | 1:52 | 28 | Unknown | 0.0291 | 0.0029 | 0.004
| 2019-07-23 | 10 | b | 64 | 1:53 | 29 | Unknown | 0.0282 | 0.0019 | 0.002
| 2019-07-23 | 10 | c | 67 | 1:54 | 30 | Unknown | 0.0289 | 0.0021 | 0.003
| 2019-07-23 | 11 | a | 68 | 1:55 | 31 | Unknown | 0.0288 | 0.0024 | 0.003
| 2019-07-23 | 11 | b | 70 | 1:56 | 32 | Unknown | 0.0284 | 0.002 | 0.003
| 2019-07-23 | 11 | c | 71 | 1:57 | 33 | Unknown | 0.0291 | 0.0027 | 0.004
| 2019-07-23 | 12 | a | 72 | 1:58 | 34 | Unknown | 0.0328 | 0.0064 | 0.008
| 2019-07-23 | 12 | b | 73 | 1:59 | 35 | Unknown | 0.0316 | 0.0051 | 0.007
| 2019-07-23 | 12 | c | 74 | 1:60 | 36 | Unknown | 0.0323 | 0.0058 | 0.007
| 2019-07-23 | 13 | a | 75 | 2:1 | 37 | Unknown | 0.0283 | 0.0017 | 0.002
| 2019-07-23 | 13 | b | 76 | 2:2 | 38 | Unknown | 0.0281 | 0.0016 | 0.002
| 2019-07-23 | 13 | c | 77 | 2:3 | 39 | Unknown | 0.0285 | 0.0019 | 0.003
| 2019-08-07 | 1 | a | 78 | 2:4 | 40 | Unknown | 0.0411 | 0.0145 | 0.019
| 2019-08-07 | 1 | b | 79 | 2:5 | 41 | Unknown | 0.041 | 0.0143 | 0.018
| 2019-08-07 | 1 | c | 81 | 2:6 | 42 | Unknown | 0.0377 | 0.0111 | 0.014
| 2019-08-07 | 2 | a | 82 | 2:7 | 43 | Unknown | 0.0437 | 0.0171 | 0.022
| 2019-08-07 | 2 | b | 83 | 2:8 | 44 | Unknown | 0.0438 | 0.017 | 0.022
| 2019-08-07 | 2 | c | 84 | 2:9 | 45 | Unknown | 0.0406 | 0.0139 | 0.018
| 2019-08-07 | 3 | a | 87 | 2:10 | 46 | Unknown | 0.0502 | 0.0231 | 0.03
| 2019-08-07 | 3 | b | 88 | 2:11 | 47 | Unknown | 0.0472 | 0.0204 | 0.026
| 2019-08-07 | 3 | c | 89 | 2:12 | 48 | Unknown | 0.0472 | 0.0205 | 0.026
| 2019-08-07 | 4 | a | 90 | 2:13 | 49 | Unknown | 0.0435 | 0.0168 | 0.022
| 2019-08-07 | 4 | b | 92 | 2:14 | 50 | Unknown | 0.0431 | 0.0164 | 0.021
| 2019-08-07 | 4 | c | 95 | 2:15 | 51 | Unknown | 0.0458 | 0.0179 | 0.023
| 2019-08-07 | 5 | a | 96 | 2:16 | 52 | Unknown | 0.0631 | 0.0359 | 0.046
| 2019-08-07 | 5 | b | 97 | 2:17 | 53 | Unknown | 0.059 | 0.0316 | 0.041
| 2019-08-07 | 5 | c | 98 | 2:18 | 54 | Unknown | 0.0701 | 0.0426 | 0.055
| 2019-08-07 | 6 | a | 99 | 2:19 | 55 | Unknown | 0.1894 | 0.1618 | 0.209
| 2019-08-07 | 6 | b | 100 | 2:20 | 56 | Unknown | 0.178 | 0.1495 | 0.193
| 2019-08-07 | 6 | c | 101 | 2:21 | 57 | Unknown | 0.1946 | 0.1661 | 0.214
| 2019-08-07 | 7 | a | 103 | 2:22 | 58 | Unknown | 0.0557 | 0.028 | 0.036
| 2019-08-07 | 7 | b | 104 | 2:23 | 59 | Unknown | 0.05 | 0.0222 | 0.029
| 2019-08-07 | 7 | c | 105 | 2:24 | 60 | Unknown | 0.0483 | 0.0206 | 0.027
| 2019-08-07 | 8 | a | 108 | 2:25 | 61 | Unknown | 0.0313 | 0.0034 | 0.004
| 2019-08-07 | 8 | b | 109 | 2:26 | 62 | Unknown | 0.0315 | 0.0041 | 0.005
| 2019-08-07 | 8 | c | 110 | 2:27 | 63 | Unknown | 0.0312 | 0.0039 | 0.005
| 2019-08-07 | 9 | a | 111 | 2:28 | 64 | Unknown | 0.0311 | 0.0038 | 0.005
| 2019-08-07 | 9 | b | 112 | 2:29 | 65 | Unknown | 0.0319 | 0.0047 | 0.006
| 2019-08-07 | 9 | c | 114 | 2:30 | 66 | Unknown | 0.0313 | 0.0041 | 0.005
| 2019-08-07 | 10 | a | 115 | 2:31 | 67 | Unknown | 0.0346 | 0.0074 | 0.01
| 2019-08-07 | 10 | b | 116 | 2:32 | 68 | Unknown | 0.0353 | 0.008 | 0.01
| 2019-08-07 | 10 | c | 117 | 2:33 | 69 | Unknown | 0.0368 | 0.0095 | 0.012
| 2019-08-07 | 11 | a | 118 | 2:34 | 70 | Unknown | 0.0316 | 0.0041 | 0.005
| 2019-08-07 | 11 | b | 119 | 2:35 | 71 | Unknown | 0.0322 | 0.0047 | 0.006
| 2019-08-07 | 11 | c | 120 | 2:36 | 72 | Unknown | 0.0325 | 0.0049 | 0.006
| 2019-08-07 | 12 | a | 121 | 2:37 | 73 | Unknown | 0.0452 | 0.0176 | 0.023
| 2019-08-07 | 12 | b | 122 | 2:38 | 74 | Unknown | 0.0462 | 0.0184 | 0.024
| 2019-08-07 | 12 | c | 123 | 2:39 | 75 | Unknown | 0.0438 | 0.016 | 0.021
| 2019-08-07 | 13 | a | 127 | 2:40 | 76 | Unknown | 0.0326 | 0.0043 | 0.006
| 2019-08-07 | 13 | b | 128 | 2:41 | 77 | Unknown | 0.0324 | 0.0044 | 0.006
| 2019-08-07 | 13 | c | 129 | 2:42 | 78 | Unknown | 0.033 | 0.005 | 0.006
| 2019-08-22 | 1 | a | 130 | 2:43 | 79 | Unknown | 0.0414 | 0.0133 | 0.017
| 2019-08-22 | 1 | b | 131 | 2:44 | 80 | Unknown | 0.0427 | 0.0145 | 0.019
| 2019-08-22 | 1 | c | 132 | 2:45 | 81 | Unknown | 0.0443 | 0.016 | 0.021
| 2019-08-22 | 2 | a | 136 | 2:46 | 82 | Unknown | 0.0395 | 0.0113 | 0.015
| 2019-08-22 | 2 | b | 137 | 2:47 | 83 | Unknown | 0.0404 | 0.0122 | 0.016
| 2019-08-22 | 2 | c | 138 | 2:48 | 84 | Unknown | 0.0408 | 0.0125 | 0.016
| 2019-08-22 | 3 | a | 139 | 2:49 | 85 | Unknown | 0.0652 | 0.037 | 0.048
| 2019-08-22 | 3 | b | 140 | 2:50 | 86 | Unknown | 0.0653 | 0.037 | 0.048
| 2019-08-22 | 3 | c | 141 | 2:51 | 87 | Unknown | 0.0653 | 0.0371 | 0.048
| 2019-08-22 | 4 | a | 142 | 2:52 | 88 | Unknown | 0.0469 | 0.0187 | 0.024
| 2019-08-22 | 4 | b | 143 | 2:53 | 89 | Unknown | 0.0461 | 0.0181 | 0.023
| 2019-08-22 | 4 | c | 147 | 2:54 | 90 | Unknown | 0.0467 | 0.0189 | 0.024
| 2019-08-22 | 5 | a | 148 | 2:55 | 91 | Unknown | 0.068 | 0.0401 | 0.052
| 2019-08-22 | 5 | b | 149 | 2:56 | 92 | Unknown | 0.0713 | 0.0433 | 0.056
| 2019-08-22 | 5 | c | 150 | 2:57 | 93 | Unknown | 0.0674 | 0.0394 | 0.051
| 2019-08-22 | 6 | a | 151 | 2:58 | 94 | Unknown | 0.1235 | 0.0955 | 0.123
| 2019-08-22 | 6 | b | 152 | 2:59 | 95 | Unknown | 0.1054 | 0.0771 | 0.099
| 2019-08-22 | 6 | c | 153 | 2:60 | 96 | Unknown | 0.1249 | 0.0967 | 0.125
| 2019-08-22 | 7 | a | 154 | 3:1 | 97 | Unknown | 0.0284 | 0 | 0
| 2019-08-22 | 7 | b | 155 | 3:2 | 98 | Unknown | 0.0991 | 0.0713 | 0.092
| 2019-08-22 | 7 | c | 156 | 3:3 | 99 | Unknown | 0.0923 | 0.0641 | 0.083
| 2019-08-22 | 8 | a | 158 | 3:4 | 100 | Unknown | 0.0919 | 0.0642 | 0.083
| 2019-08-22 | 8 | b | 159 | 3:5 | 101 | Unknown | 0.0884 | 0.0603 | 0.078
| 2019-08-22 | 8 | c | 160 | 3:6 | 102 | Unknown | 0.0642 | 0.0361 | 0.047
| 2019-08-22 | 9 | a | 161 | 3:7 | 103 | Unknown | 0.0326 | 0.0047 | 0.006
| 2019-08-22 | 9 | b | 162 | 3:8 | 104 | Unknown | 0.0324 | 0.0046 | 0.006
| 2019-08-22 | 9 | c | 163 | 3:9 | 105 | Unknown | 0.0326 | 0.0049 | 0.006
| 2019-08-22 | 10 | a | 164 | 3:10 | 106 | Unknown | 0.0334 | 0.0057 | 0.007
| 2019-08-22 | 10 | b | 165 | 3:11 | 107 | Unknown | 0.0319 | 0.0041 | 0.005
| 2019-08-22 | 10 | c | 166 | 3:12 | 108 | Unknown | 0.0333 | 0.0056 | 0.007
| 2019-08-22 | 11 | a | 167 | 3:13 | 109 | Unknown | 0.0309 | 0.0032 | 0.004
| 2019-08-22 | 11 | b | 171 | 3:14 | 110 | Unknown | 0.031 | 0.0027 | 0.003
| 2019-08-22 | 11 | c | 172 | 3:15 | 111 | Unknown | 0.0309 | 0.003 | 0.004
| 2019-08-22 | 12 | a | 173 | 3:16 | 112 | Unknown | 0.0347 | 0.0068 | 0.009
| 2019-08-22 | 12 | b | 174 | 3:17 | 113 | Unknown | 0.0335 | 0.0055 | 0.007
| 2019-08-22 | 12 | c | 175 | 3:18 | 114 | Unknown | 0.0357 | 0.0076 | 0.01
| 2019-08-22 | 13 | a | 176 | 3:19 | 115 | Unknown | 0.0302 | 0.0021 | 0.003
| 2019-08-22 | 13 | b | 177 | 3:20 | 116 | Unknown | 0.0298 | 0.0017 | 0.002
| 2019-08-22 | 13 | c | 178 | 3:21 | 117 | Unknown | 0.0301 | 0.0019 | 0.002
| LB8/7/2019 |  | a | 184 | 3:22 | 118 | Unknown | 0.0337 | 0.0047 | 0.006
| LB8/7/2019 |  | b | 185 | 3:23 | 119 | Unknown | 0.0331 | 0.0051 | 0.007
| LB8/7/2019 |  | c | 186 | 3:24 | 120 | Unknown | 0.0334 | 0.0055 | 0.007
| JB8/7/2019 |  | a | 187 | 3:25 | 121 | Unknown | 0.0304 | 0.0025 | 0.003
| JB8/7/2019 |  | b | 188 | 3:26 | 122 | Unknown | 0.0303 | 0.0025 | 0.003
| JB8/7/2019 |  | c | 189 | 3:27 | 123 | Unknown | 0.0305 | 0.0027 | 0.004
| LB8/22/2019 |  | a | 191 | 3:28 | 124 | Unknown | 0.0367 | 0.009 | 0.012
| LB8/22/2019 |  | b | 192 | 3:29 | 125 | Unknown | 0.0351 | 0.0073 | 0.009
| LB8/22/2019 |  | c | 195 | 3:30 | 126 | Unknown | 0.0345 | 0.0062 | 0.008
| JB8/22/2019 |  | a | 196 | 3:31 | 127 | Unknown | 0.0319 | 0.004 | 0.005
| JB8/22/2019 |  | b | 197 | 3:32 | 128 | Unknown | 0.031 | 0.0031 | 0.004
| JB8/22/2019 |  | c | 198 | 3:33 | 129 | Unknown | 0.032 | 0.0041 | 0.005

## Std Curve Construction
1. Copy in Calibrant Information
[10 empty] | 2. Run Linear Regression | [8 empty] | 4. If quadratic exceeds 10%, run split-curve procedure.
Use X=Peak height; Y=Conc)
| Calibrant |  | Linear Fit |  | Quadratic |  | Split Curve
Calibrant | Conc. | Calibrant | Std Conc |  | Std Conc |  | Std Conc | [11 empty] | Plot low range calibrant data: |  |  | Enter coefficients below graph.
Id | (mg/L N) | peak hght (x) | (y-hat) | % error | (y-hat) | % error | (y-hat) | % error |  |  | All data - Linear Fit
C1 | 0.001 | 0.0006 | 0.00078165 | -21.8345 | 0.0010657 | 6.5732 | 0.00141 | 40.9966 |  | SUMMARY OUTPUT
C1 | 0.001 | 0.0002 | 0.00026585 | -73.4154 | 0.00055524 | -44.4759
C1 | 0.001 | -0.0002 | -0.00024996 | -124.9963 | 4.4761e-05 | -95.5239 |  |  |  | Regression Statistics
C2 | 0.005 | 0.0042 | 0.0054239 | 8.4787 | 0.0056606 | 13.2125 |  |  |  | Multiple R | 0.99999
C2 | 0.005 | 0.0038 | 0.0049081 | -1.8375 | 0.00515 | 3.0008 |  |  |  | R Square | 0.99998
C2 | 0.005 | 0.0036 | 0.0046502 | -6.9956 | 0.0048948 | -2.105 |  |  |  | Adjusted R Square | 0.99998
C3 | 0.01 | 0.0071 | 0.0091636 | -8.3645 | 0.0093627 | -6.373 |  |  |  | Standard Error | 0.00087232
C3 | 0.01 | 0.0081 | 0.010453 | 4.5307 | 0.010639 | 6.394 |  |  |  | Observations | 18
C3 | 0.01 | 0.0077 | 0.0099373 | -0.62736 | 0.010129 | 1.2871 |  |  | Split
C4 | 0.05 | 0.0397 | 0.051202 | 2.404 | 0.051018 | 2.0352 |  |  |  | ANOVA
C4 | 0.05 | 0.0391 | 0.050428 | 0.85653 | 0.05025 | 0.50066 |  |  |  |  | df | SS | MS | F | Significance F
C4 | 0.05 | 0.0393 | 0.050686 | 1.3723 | 0.050506 | 1.0122 |  |  |  | Regression | 1 | 0.56609 | 0.56609 | 743926.4168 | 8.9898e-39
C5 | 0.1 | 0.0775 | 0.099946 | -0.054073 | 0.099406 | -0.59449 |  |  |  | Residual | 16 | 1.2175e-05 | 7.6095e-07
C5 | 0.1 | 0.0776 | 0.10007 | 0.07488 | 0.099534 | -0.46636 |  |  |  | Total | 17 | 0.5661
C5 | 0.1 | 0.0781 | 0.10072 | 0.71964 | 0.10017 | 0.17434
C6 | 0.5 | 0.3892 | 0.50189 | 0.37801 | 0.50204 | 0.40824 |  |  |  |  | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
C6 | 0.5 | 0.3869 | 0.49892 | -0.21517 | 0.49905 | -0.1907 |  |  |  | Intercept | 7.9415e-06 | 0.00024256 | 0.032741 | 0.97429 | -0.00050626 | 0.00052214 | -0.00050626 | 0.00052214 | Low curve coefficients:
C6 | 0.5 | 0.3868 | 0.4988 | -0.24096 | 0.49892 | -0.21674 |  |  |  | X Variable 1 | 1.2895 | 0.0014951 | 862.5117 | 8.9898e-39 | 1.2864 | 1.2927 | 1.2864 | 1.2927 |  | intcpt | X coef. | X^2 coef.
[20 empty] | 0.0007 | 1.1822 | 1.7933
[10 empty] | (Use X=Peak Ht.; Y=Conc) | [8 empty] | Plot High Range calibrant data. |  |  | Enter coefficients below graph.
Enter equation parameters below graph.
[10 empty] | Quadratic curve coefficients | [8 empty] | High curve coefficients:
[13 empty] | Intercept | X coef. | X^2 coef. |  |  |  |  | intcpt | X coef | X^2 coef
[13 empty] | 0.0003 | 1.2762 | 0.0333 |  |  |  |  | -0.0015 | 1.31 | -0.0418
Show "Split" in spreadsheet.

## Apply Std Curve
|  |  | Sample Info |  |  |  | FASPacN-NO3 |  |  | Curve-fit results N-NO3
SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | Peak# | Position | Identifier | Type | Raw Ht | Cor Ht | mg/L |  | mg/L |  | Curve-fitting parameters used:
2019-07-23 | 1 | a | 31 | 1:25 | 1 | Unknown | 0.0485 | 0.0223 | 0.029 |  | 0.027955 | Copy from Std Curve Construction page.
2019-07-23 | 1 | b | 32 | 1:26 | 2 | Unknown | 0.0449 | 0.0186 | 0.024 |  | 0.023309
2019-07-23 | 1 | c | 33 | 1:27 | 3 | Unknown | 0.0492 | 0.023 | 0.03 |  | 0.028839 | Use these for linear or quadratic fit.
2019-07-23 | 2 | a | 34 | 1:28 | 4 | Unknown | 0.0444 | 0.0183 | 0.024 |  | 0.022935 | If linear, enter 0 for X^2 coefficient.
2019-07-23 | 2 | b | 35 | 1:29 | 5 | Unknown | 0.0432 | 0.0172 | 0.022 |  | 0.021564
2019-07-23 | 2 | c | 37 | 1:30 | 6 | Unknown | 0.0439 | 0.0181 | 0.023 |  | 0.022685 | Intercept: | X Coef. | X^2 Coef:
2019-07-23 | 3 | a | 38 | 1:31 | 7 | Unknown | 0.0415 | 0.0156 | 0.02 |  | 0.019579 | -0.0267 | 1.1963 | 0.1673
2019-07-23 | 3 | b | 39 | 1:32 | 8 | Unknown | 0.043 | 0.017 | 0.022 |  | 0.021316
2019-07-23 | 3 | c | 40 | 1:33 | 9 | Unknown | 0.0431 | 0.0171 | 0.022 |  | 0.02144
2019-07-23 | 4 | a | 41 | 1:34 | 10 | Unknown | 0.0353 | 0.0093 | 0.012 |  | 0.01185 | Use these for split curve:
2019-07-23 | 4 | b | 42 | 1:35 | 11 | Unknown | 0.0354 | 0.0095 | 0.012 |  | 0.012093 | Low Range:
2019-07-23 | 4 | c | 43 | 1:36 | 12 | Unknown | 0.0345 | 0.0085 | 0.011 |  | 0.010878 | Intercept: | X Coef. | X^2 Coef:
2019-07-23 | 5 | a | 44 | 1:37 | 13 | Unknown | 0.0352 | 0.0092 | 0.012 |  | 0.011728 | 0.0007 | 1.1822 | 1.7933
2019-07-23 | 5 | b | 45 | 1:38 | 14 | Unknown | 0.0352 | 0.0092 | 0.012 |  | 0.011728
2019-07-23 | 5 | c | 46 | 1:39 | 15 | Unknown | 0.0352 | 0.0092 | 0.012 |  | 0.011728
2019-07-23 | 6 | a | 50 | 1:40 | 16 | Unknown | 0.0445 | 0.0179 | 0.023 |  | 0.022436 | High Range:
2019-07-23 | 6 | b | 51 | 1:41 | 17 | Unknown | 0.0431 | 0.017 | 0.022 |  | 0.021316 | Intercept: | X Coef. | X^2 Coef:
2019-07-23 | 6 | c | 52 | 1:42 | 18 | Unknown | 0.0435 | 0.0173 | 0.022 |  | 0.021689 | -0.0015 | 1.31 | -0.0418
2019-07-23 | 7 | a | 53 | 1:43 | 19 | Unknown | 0.0288 | 0.0026 | 0.003 |  | 0.0037858
2019-07-23 | 7 | b | 54 | 1:44 | 20 | Unknown | 0.0284 | 0.0022 | 0.003 |  | 0.0033095 | Procedure for split curve:
2019-07-23 | 7 | c | 55 | 1:45 | 21 | Unknown | 0.0292 | 0.0031 | 0.004 |  | 0.0043821 | 1. Sort data on Row I: FASpac results, low to high.
2019-07-23 | 8 | a | 56 | 1:46 | 22 | Unknown | 0.029 | 0.0029 | 0.004 |  | 0.0041435 | 2. Apply low range equation.
2019-07-23 | 8 | b | 57 | 1:47 | 23 | Unknown | 0.0284 | 0.0023 | 0.003 |  | 0.0034285 | 3. Apply high range equation.
2019-07-23 | 8 | c | 59 | 1:48 | 24 | Unknown | 0.0281 | 0.0019 | 0.003 |  | 0.0029527 | 4. Note here peak height where split occurs: |  |  | 0.05
2019-07-23 | 9 | a | 60 | 1:49 | 25 | Unknown | 0.0278 | 0.0016 | 0.002 |  | 0.0025961 | 4. Sort back based on Row D: Peak Number
2019-07-23 | 9 | b | 61 | 1:50 | 26 | Unknown | 0.0277 | 0.0014 | 0.002 |  | 0.0023586
2019-07-23 | 9 | c | 62 | 1:51 | 27 | Unknown | 0.028 | 0.0017 | 0.002 |  | 0.0027149
2019-07-23 | 10 | a | 63 | 1:52 | 28 | Unknown | 0.0291 | 0.0029 | 0.004 |  | 0.0041435
2019-07-23 | 10 | b | 64 | 1:53 | 29 | Unknown | 0.0282 | 0.0019 | 0.002 |  | 0.0029527
2019-07-23 | 10 | c | 67 | 1:54 | 30 | Unknown | 0.0289 | 0.0021 | 0.003 |  | 0.0031905
2019-07-23 | 11 | a | 68 | 1:55 | 31 | Unknown | 0.0288 | 0.0024 | 0.003 |  | 0.0035476
2019-07-23 | 11 | b | 70 | 1:56 | 32 | Unknown | 0.0284 | 0.002 | 0.003 |  | 0.0030716
2019-07-23 | 11 | c | 71 | 1:57 | 33 | Unknown | 0.0291 | 0.0027 | 0.004 |  | 0.003905
2019-07-23 | 12 | a | 72 | 1:58 | 34 | Unknown | 0.0328 | 0.0064 | 0.008 |  | 0.0083395
2019-07-23 | 12 | b | 73 | 1:59 | 35 | Unknown | 0.0316 | 0.0051 | 0.007 |  | 0.0067759
2019-07-23 | 12 | c | 74 | 1:60 | 36 | Unknown | 0.0323 | 0.0058 | 0.007 |  | 0.0076171
2019-07-23 | 13 | a | 75 | 2:1 | 37 | Unknown | 0.0283 | 0.0017 | 0.002 |  | 0.0027149
2019-07-23 | 13 | b | 76 | 2:2 | 38 | Unknown | 0.0281 | 0.0016 | 0.002 |  | 0.0025961
2019-07-23 | 13 | c | 77 | 2:3 | 39 | Unknown | 0.0285 | 0.0019 | 0.003 |  | 0.0029527
2019-08-07 | 1 | a | 78 | 2:4 | 40 | Unknown | 0.0411 | 0.0145 | 0.019 |  | 0.018219
2019-08-07 | 1 | b | 79 | 2:5 | 41 | Unknown | 0.041 | 0.0143 | 0.018 |  | 0.017972
2019-08-07 | 1 | c | 81 | 2:6 | 42 | Unknown | 0.0377 | 0.0111 | 0.014 |  | 0.014043
2019-08-07 | 2 | a | 82 | 2:7 | 43 | Unknown | 0.0437 | 0.0171 | 0.022 |  | 0.02144
2019-08-07 | 2 | b | 83 | 2:8 | 44 | Unknown | 0.0438 | 0.017 | 0.022 |  | 0.021316
2019-08-07 | 2 | c | 84 | 2:9 | 45 | Unknown | 0.0406 | 0.0139 | 0.018 |  | 0.017479
2019-08-07 | 3 | a | 87 | 2:10 | 46 | Unknown | 0.0502 | 0.0231 | 0.03 |  | 0.028966
2019-08-07 | 3 | b | 88 | 2:11 | 47 | Unknown | 0.0472 | 0.0204 | 0.026 |  | 0.025563
2019-08-07 | 3 | c | 89 | 2:12 | 48 | Unknown | 0.0472 | 0.0205 | 0.026 |  | 0.025689
2019-08-07 | 4 | a | 90 | 2:13 | 49 | Unknown | 0.0435 | 0.0168 | 0.022 |  | 0.021067
2019-08-07 | 4 | b | 92 | 2:14 | 50 | Unknown | 0.0431 | 0.0164 | 0.021 |  | 0.02057
2019-08-07 | 4 | c | 95 | 2:15 | 51 | Unknown | 0.0458 | 0.0179 | 0.023 |  | 0.022436
2019-08-07 | 5 | a | 96 | 2:16 | 52 | Unknown | 0.0631 | 0.0359 | 0.046 |  | 0.045452
2019-08-07 | 5 | b | 97 | 2:17 | 53 | Unknown | 0.059 | 0.0316 | 0.041 |  | 0.039848
2019-08-07 | 5 | c | 98 | 2:18 | 54 | Unknown | 0.0701 | 0.0426 | 0.055 |  | 0.05423
2019-08-07 | 6 | a | 99 | 2:19 | 55 | Unknown | 0.1894 | 0.1618 | 0.209 |  | 0.20936
2019-08-07 | 6 | b | 100 | 2:20 | 56 | Unknown | 0.178 | 0.1495 | 0.193 |  | 0.19341
2019-08-07 | 6 | c | 101 | 2:21 | 57 | Unknown | 0.1946 | 0.1661 | 0.214 |  | 0.21494
2019-08-07 | 7 | a | 103 | 2:22 | 58 | Unknown | 0.0557 | 0.028 | 0.036 |  | 0.035208
2019-08-07 | 7 | b | 104 | 2:23 | 59 | Unknown | 0.05 | 0.0222 | 0.029 |  | 0.027829
2019-08-07 | 7 | c | 105 | 2:24 | 60 | Unknown | 0.0483 | 0.0206 | 0.027 |  | 0.025814
2019-08-07 | 8 | a | 108 | 2:25 | 61 | Unknown | 0.0313 | 0.0034 | 0.004 |  | 0.0047402
2019-08-07 | 8 | b | 109 | 2:26 | 62 | Unknown | 0.0315 | 0.0041 | 0.005 |  | 0.0055772
2019-08-07 | 8 | c | 110 | 2:27 | 63 | Unknown | 0.0312 | 0.0039 | 0.005 |  | 0.0053379
2019-08-07 | 9 | a | 111 | 2:28 | 64 | Unknown | 0.0311 | 0.0038 | 0.005 |  | 0.0052183
2019-08-07 | 9 | b | 112 | 2:29 | 65 | Unknown | 0.0319 | 0.0047 | 0.006 |  | 0.006296
2019-08-07 | 9 | c | 114 | 2:30 | 66 | Unknown | 0.0313 | 0.0041 | 0.005 |  | 0.0055772
2019-08-07 | 10 | a | 115 | 2:31 | 67 | Unknown | 0.0346 | 0.0074 | 0.01 |  | 0.0095465
2019-08-07 | 10 | b | 116 | 2:32 | 68 | Unknown | 0.0353 | 0.008 | 0.01 |  | 0.010272
2019-08-07 | 10 | c | 117 | 2:33 | 69 | Unknown | 0.0368 | 0.0095 | 0.012 |  | 0.012093
2019-08-07 | 11 | a | 118 | 2:34 | 70 | Unknown | 0.0316 | 0.0041 | 0.005 |  | 0.0055772
2019-08-07 | 11 | b | 119 | 2:35 | 71 | Unknown | 0.0322 | 0.0047 | 0.006 |  | 0.006296
2019-08-07 | 11 | c | 120 | 2:36 | 72 | Unknown | 0.0325 | 0.0049 | 0.006 |  | 0.0065358
2019-08-07 | 12 | a | 121 | 2:37 | 73 | Unknown | 0.0452 | 0.0176 | 0.023 |  | 0.022062
2019-08-07 | 12 | b | 122 | 2:38 | 74 | Unknown | 0.0462 | 0.0184 | 0.024 |  | 0.02306
2019-08-07 | 12 | c | 123 | 2:39 | 75 | Unknown | 0.0438 | 0.016 | 0.021 |  | 0.020074
2019-08-07 | 13 | a | 127 | 2:40 | 76 | Unknown | 0.0326 | 0.0043 | 0.006 |  | 0.0058166
2019-08-07 | 13 | b | 128 | 2:41 | 77 | Unknown | 0.0324 | 0.0044 | 0.006 |  | 0.0059364
2019-08-07 | 13 | c | 129 | 2:42 | 78 | Unknown | 0.033 | 0.005 | 0.006 |  | 0.0066558
2019-08-22 | 1 | a | 130 | 2:43 | 79 | Unknown | 0.0414 | 0.0133 | 0.017 |  | 0.01674
2019-08-22 | 1 | b | 131 | 2:44 | 80 | Unknown | 0.0427 | 0.0145 | 0.019 |  | 0.018219
2019-08-22 | 1 | c | 132 | 2:45 | 81 | Unknown | 0.0443 | 0.016 | 0.021 |  | 0.020074
2019-08-22 | 2 | a | 136 | 2:46 | 82 | Unknown | 0.0395 | 0.0113 | 0.015 |  | 0.014288
2019-08-22 | 2 | b | 137 | 2:47 | 83 | Unknown | 0.0404 | 0.0122 | 0.016 |  | 0.01539
2019-08-22 | 2 | c | 138 | 2:48 | 84 | Unknown | 0.0408 | 0.0125 | 0.016 |  | 0.015758
2019-08-22 | 3 | a | 139 | 2:49 | 85 | Unknown | 0.0652 | 0.037 | 0.048 |  | 0.046896
2019-08-22 | 3 | b | 140 | 2:50 | 86 | Unknown | 0.0653 | 0.037 | 0.048 |  | 0.046896
2019-08-22 | 3 | c | 141 | 2:51 | 87 | Unknown | 0.0653 | 0.0371 | 0.048 |  | 0.047028
2019-08-22 | 4 | a | 142 | 2:52 | 88 | Unknown | 0.0469 | 0.0187 | 0.024 |  | 0.023434
2019-08-22 | 4 | b | 143 | 2:53 | 89 | Unknown | 0.0461 | 0.0181 | 0.023 |  | 0.022685
2019-08-22 | 4 | c | 147 | 2:54 | 90 | Unknown | 0.0467 | 0.0189 | 0.024 |  | 0.023684
2019-08-22 | 5 | a | 148 | 2:55 | 91 | Unknown | 0.068 | 0.0401 | 0.052 |  | 0.050964
2019-08-22 | 5 | b | 149 | 2:56 | 92 | Unknown | 0.0713 | 0.0433 | 0.056 |  | 0.055145
2019-08-22 | 5 | c | 150 | 2:57 | 93 | Unknown | 0.0674 | 0.0394 | 0.051 |  | 0.050049
2019-08-22 | 6 | a | 151 | 2:58 | 94 | Unknown | 0.1235 | 0.0955 | 0.123 |  | 0.12322
2019-08-22 | 6 | b | 152 | 2:59 | 95 | Unknown | 0.1054 | 0.0771 | 0.099 |  | 0.099253
2019-08-22 | 6 | c | 153 | 2:60 | 96 | Unknown | 0.1249 | 0.0967 | 0.125 |  | 0.12479
2019-08-22 | 7 | a | 154 | 3:1 | 97 | Unknown | 0.0284 | 0 | 0 |  | 0.0007
2019-08-22 | 7 | b | 155 | 3:2 | 98 | Unknown | 0.0991 | 0.0713 | 0.092 |  | 0.091691
2019-08-22 | 7 | c | 156 | 3:3 | 99 | Unknown | 0.0923 | 0.0641 | 0.083 |  | 0.082299
2019-08-22 | 8 | a | 158 | 3:4 | 100 | Unknown | 0.0919 | 0.0642 | 0.083 |  | 0.08243
2019-08-22 | 8 | b | 159 | 3:5 | 101 | Unknown | 0.0884 | 0.0603 | 0.078 |  | 0.077341
2019-08-22 | 8 | c | 160 | 3:6 | 102 | Unknown | 0.0642 | 0.0361 | 0.047 |  | 0.045714
2019-08-22 | 9 | a | 161 | 3:7 | 103 | Unknown | 0.0326 | 0.0047 | 0.006 |  | 0.006296
2019-08-22 | 9 | b | 162 | 3:8 | 104 | Unknown | 0.0324 | 0.0046 | 0.006 |  | 0.0061761
2019-08-22 | 9 | c | 163 | 3:9 | 105 | Unknown | 0.0326 | 0.0049 | 0.006 |  | 0.0065358
2019-08-22 | 10 | a | 164 | 3:10 | 106 | Unknown | 0.0334 | 0.0057 | 0.007 |  | 0.0074968
2019-08-22 | 10 | b | 165 | 3:11 | 107 | Unknown | 0.0319 | 0.0041 | 0.005 |  | 0.0055772
2019-08-22 | 10 | c | 166 | 3:12 | 108 | Unknown | 0.0333 | 0.0056 | 0.007 |  | 0.0073766
2019-08-22 | 11 | a | 167 | 3:13 | 109 | Unknown | 0.0309 | 0.0032 | 0.004 |  | 0.0045014
2019-08-22 | 11 | b | 171 | 3:14 | 110 | Unknown | 0.031 | 0.0027 | 0.003 |  | 0.003905
2019-08-22 | 11 | c | 172 | 3:15 | 111 | Unknown | 0.0309 | 0.003 | 0.004 |  | 0.0042627
2019-08-22 | 12 | a | 173 | 3:16 | 112 | Unknown | 0.0347 | 0.0068 | 0.009 |  | 0.0088219
2019-08-22 | 12 | b | 174 | 3:17 | 113 | Unknown | 0.0335 | 0.0055 | 0.007 |  | 0.0072563
2019-08-22 | 12 | c | 175 | 3:18 | 114 | Unknown | 0.0357 | 0.0076 | 0.01 |  | 0.0097883
2019-08-22 | 13 | a | 176 | 3:19 | 115 | Unknown | 0.0302 | 0.0021 | 0.003 |  | 0.0031905
2019-08-22 | 13 | b | 177 | 3:20 | 116 | Unknown | 0.0298 | 0.0017 | 0.002 |  | 0.0027149
2019-08-22 | 13 | c | 178 | 3:21 | 117 | Unknown | 0.0301 | 0.0019 | 0.002 |  | 0.0029527
LB8/7/2019 |  | a | 184 | 3:22 | 118 | Unknown | 0.0337 | 0.0047 | 0.006 |  | 0.006296
LB8/7/2019 |  | b | 185 | 3:23 | 119 | Unknown | 0.0331 | 0.0051 | 0.007 |  | 0.0067759
LB8/7/2019 |  | c | 186 | 3:24 | 120 | Unknown | 0.0334 | 0.0055 | 0.007 |  | 0.0072563
JB8/7/2019 |  | a | 187 | 3:25 | 121 | Unknown | 0.0304 | 0.0025 | 0.003 |  | 0.0036667
JB8/7/2019 |  | b | 188 | 3:26 | 122 | Unknown | 0.0303 | 0.0025 | 0.003 |  | 0.0036667
JB8/7/2019 |  | c | 189 | 3:27 | 123 | Unknown | 0.0305 | 0.0027 | 0.004 |  | 0.003905
LB8/22/2019 |  | a | 191 | 3:28 | 124 | Unknown | 0.0367 | 0.009 | 0.012 |  | 0.011485
LB8/22/2019 |  | b | 192 | 3:29 | 125 | Unknown | 0.0351 | 0.0073 | 0.009 |  | 0.0094256
LB8/22/2019 |  | c | 195 | 3:30 | 126 | Unknown | 0.0345 | 0.0062 | 0.008 |  | 0.0080986
JB8/22/2019 |  | a | 196 | 3:31 | 127 | Unknown | 0.0319 | 0.004 | 0.005 |  | 0.0054575
JB8/22/2019 |  | b | 197 | 3:32 | 128 | Unknown | 0.031 | 0.0031 | 0.004 |  | 0.0043821
JB8/22/2019 |  | c | 198 | 3:33 | 129 | Unknown | 0.032 | 0.0041 | 0.005 |  | 0.0055772

## Final Results
SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | N-NO3 mg/L
2019-07-23 | 1 | a | 0.027955
2019-07-23 | 1 | b | 0.023309
2019-07-23 | 1 | c | 0.028839
2019-07-23 | 2 | a | 0.022935
2019-07-23 | 2 | b | 0.021564
2019-07-23 | 2 | c | 0.022685
2019-07-23 | 3 | a | 0.019579
2019-07-23 | 3 | b | 0.021316
2019-07-23 | 3 | c | 0.02144
2019-07-23 | 4 | a | 0.01185
2019-07-23 | 4 | b | 0.012093
2019-07-23 | 4 | c | 0.010878
2019-07-23 | 5 | a | 0.011728
2019-07-23 | 5 | b | 0.011728
2019-07-23 | 5 | c | 0.011728
2019-07-23 | 6 | a | 0.022436
2019-07-23 | 6 | b | 0.021316
2019-07-23 | 6 | c | 0.021689
2019-07-23 | 7 | a | 0.0037858
2019-07-23 | 7 | b | 0.0033095
2019-07-23 | 7 | c | 0.0043821
2019-07-23 | 8 | a | 0.0041435
2019-07-23 | 8 | b | 0.0034285
2019-07-23 | 8 | c | 0.0029527
2019-07-23 | 9 | a | 0.0025961
2019-07-23 | 9 | b | 0.0023586
2019-07-23 | 9 | c | 0.0027149
2019-07-23 | 10 | a | 0.0041435
2019-07-23 | 10 | b | 0.0029527
2019-07-23 | 10 | c | 0.0031905
2019-07-23 | 11 | a | 0.0035476
2019-07-23 | 11 | b | 0.0030716
2019-07-23 | 11 | c | 0.003905
2019-07-23 | 12 | a | 0.0083395
2019-07-23 | 12 | b | 0.0067759
2019-07-23 | 12 | c | 0.0076171
2019-07-23 | 13 | a | 0.0027149
2019-07-23 | 13 | b | 0.0025961
2019-07-23 | 13 | c | 0.0029527
2019-08-07 | 1 | a | 0.018219
2019-08-07 | 1 | b | 0.017972
2019-08-07 | 1 | c | 0.014043
2019-08-07 | 2 | a | 0.02144
2019-08-07 | 2 | b | 0.021316
2019-08-07 | 2 | c | 0.017479
2019-08-07 | 3 | a | 0.028966
2019-08-07 | 3 | b | 0.025563
2019-08-07 | 3 | c | 0.025689
2019-08-07 | 4 | a | 0.021067
2019-08-07 | 4 | b | 0.02057
2019-08-07 | 4 | c | 0.022436
2019-08-07 | 5 | a | 0.045452
2019-08-07 | 5 | b | 0.039848
2019-08-07 | 5 | c | 0.05423
2019-08-07 | 6 | a | 0.20936
2019-08-07 | 6 | b | 0.19341
2019-08-07 | 6 | c | 0.21494
2019-08-07 | 7 | a | 0.035208
2019-08-07 | 7 | b | 0.027829
2019-08-07 | 7 | c | 0.025814
2019-08-07 | 8 | a | 0.0047402
2019-08-07 | 8 | b | 0.0055772
2019-08-07 | 8 | c | 0.0053379
2019-08-07 | 9 | a | 0.0052183
2019-08-07 | 9 | b | 0.006296
2019-08-07 | 9 | c | 0.0055772
2019-08-07 | 10 | a | 0.0095465
2019-08-07 | 10 | b | 0.010272
2019-08-07 | 10 | c | 0.012093
2019-08-07 | 11 | a | 0.0055772
2019-08-07 | 11 | b | 0.006296
2019-08-07 | 11 | c | 0.0065358
2019-08-07 | 12 | a | 0.022062
2019-08-07 | 12 | b | 0.02306
2019-08-07 | 12 | c | 0.020074
2019-08-07 | 13 | a | 0.0058166
2019-08-07 | 13 | b | 0.0059364
2019-08-07 | 13 | c | 0.0066558
2019-08-22 | 1 | a | 0.01674
2019-08-22 | 1 | b | 0.018219
2019-08-22 | 1 | c | 0.020074
2019-08-22 | 2 | a | 0.014288
2019-08-22 | 2 | b | 0.01539
2019-08-22 | 2 | c | 0.015758
2019-08-22 | 3 | a | 0.046896
2019-08-22 | 3 | b | 0.046896
2019-08-22 | 3 | c | 0.047028
2019-08-22 | 4 | a | 0.023434
2019-08-22 | 4 | b | 0.022685
2019-08-22 | 4 | c | 0.023684
2019-08-22 | 5 | a | 0.050964
2019-08-22 | 5 | b | 0.055145
2019-08-22 | 5 | c | 0.050049
2019-08-22 | 6 | a | 0.12322
2019-08-22 | 6 | b | 0.099253
2019-08-22 | 6 | c | 0.12479
2019-08-22 | 7 | a | 0.0007
2019-08-22 | 7 | b | 0.091691
2019-08-22 | 7 | c | 0.082299
2019-08-22 | 8 | a | 0.08243
2019-08-22 | 8 | b | 0.077341
2019-08-22 | 8 | c | 0.045714
2019-08-22 | 9 | a | 0.006296
2019-08-22 | 9 | b | 0.0061761
2019-08-22 | 9 | c | 0.0065358
2019-08-22 | 10 | a | 0.0074968
2019-08-22 | 10 | b | 0.0055772
2019-08-22 | 10 | c | 0.0073766
2019-08-22 | 11 | a | 0.0045014
2019-08-22 | 11 | b | 0.003905
2019-08-22 | 11 | c | 0.0042627
2019-08-22 | 12 | a | 0.0088219
2019-08-22 | 12 | b | 0.0072563
2019-08-22 | 12 | c | 0.0097883
2019-08-22 | 13 | a | 0.0031905
2019-08-22 | 13 | b | 0.0027149
2019-08-22 | 13 | c | 0.0029527
LB8/7/2019 | 8am | a | 0.006296
LB8/7/2019 | 8am | b | 0.0067759
LB8/7/2019 | 8am | c | 0.0072563
JB8/7/2019 | 9a | a | 0.0036667
JB8/7/2019 | 9a | b | 0.0036667
JB8/7/2019 | 9a | c | 0.003905
LB8/22/2019 | 8am | a | 0.011485
LB8/22/2019 | 8am | b | 0.0094256
LB8/22/2019 | 8am | c | 0.0080986
JB8/22/2019 | 9a | a | 0.0054575
JB8/22/2019 | 9a | b | 0.0043821
JB8/22/2019 | 9a | c | 0.0055772