## Results and Sample ID
Run Name: LTREB_NO3_AP2_2019_12_12
Run Date: 12/12/2019
Operator: Valett Lab
|  |  |  | Sample Info |  |  |  | FASPac NO3-NO2 |  |  | Calibrant | Conc., mg/L
Row | SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | Peak# | Position | Identifier | Type | Raw Ht | Cor Ht | mg/L | C1 | 0.001
|  |  |  | 1 | 1:1 | SYNC | SYNC | 0.373 | 0.3728 | 0.5 | C2 | 0.005
|  |  |  | 2 | 1:2 | CO | Carry over | 0.0029 | 0.0027 | 0.006 | C3 | 0.01
|  |  |  | 3 | 1:3 | W | Wash | 0.0002 | 0 | 0.002 | C4 | 0.05
|  |  |  | 4 | 1:4 | NO3 | Unknown | 0.0742 | 0.0739 | 0.101 | C5 | 0.1
|  |  |  | 5 | 1:5 | NOX% | NO3 check | 0.078 | 0.077 | 0.105 | C6 | 0.5
[6 empty] | NO3 Efficiency |  |  |  | 96.09%
|  |  |  | 6 | 1:3 | W | Wash | 0.0006 | 0 | 0.002
|  |  |  | 7 | 1:7 | C1 | Calibrant | 0.0009 | 0.0002 | 0.002
|  |  |  | 8 | 1:7 | C1 | Calibrant | 0.0009 | 0 | 0.002
|  |  |  | 9 | 1:7 | C1 | Calibrant | 0.0009 | -0.0001 | 0.002
|  |  |  | 10 | 1:8 | C2 | Calibrant | 0.0038 | 0.0026 | 0.006
|  |  |  | 11 | 1:8 | C2 | Calibrant | 0.0037 | 0.0024 | 0.005
|  |  |  | 12 | 1:8 | C2 | Calibrant | 0.0039 | 0.0024 | 0.005
|  |  |  | 13 | 1:9 | C3 | Calibrant | 0.0076 | 0.006 | 0.01
|  |  |  | 14 | 1:9 | C3 | Calibrant | 0.0075 | 0.0058 | 0.01
|  |  |  | 15 | 1:9 | C3 | Calibrant | 0.0076 | 0.0057 | 0.01
|  |  |  | 16 | 1:10 | C4 | Calibrant | 0.0372 | 0.0352 | 0.049
|  |  |  | 17 | 1:10 | C4 | Calibrant | 0.0378 | 0.0354 | 0.049
|  |  |  | 18 | 1:10 | C4 | Calibrant | 0.0382 | 0.0356 | 0.05
|  |  |  | 19 | 1:11 | C5 | Calibrant | 0.0748 | 0.0721 | 0.098
|  |  |  | 20 | 1:11 | C5 | Calibrant | 0.0754 | 0.0723 | 0.099
|  |  |  | 21 | 1:11 | C5 | Calibrant | 0.0754 | 0.0721 | 0.099
|  |  |  | 22 | 1:12 | C6 | Calibrant | 0.3767 | 0.3733 | 0.501
|  |  |  | 23 | 1:12 | C6 | Calibrant | 0.3799 | 0.3742 | 0.502
|  |  |  | 24 | 1:12 | C6 | Calibrant | 0.3774 | 0.3716 | 0.498
|  |  |  | 25 | 1:3 | W | Wash | 0.0032 | 0 | 0.002
|  |  |  | 26 | 1:4 | NO3 | Unknown | 0.0746 | 0.0718 | 0.098
|  |  |  | 27 | 1:5 | NOX% | NO3 check | 0.0778 | 0.0748 | 0.102
[6 empty] | NO3 Efficiency |  |  |  | 95.99%
|  |  |  | 28 | 1:13 | CC1 | QC 2 | 0.016 | 0.0134 | 0.02
|  |  |  | 29 | 1:14 | CC2 | Unknown | 0.1484 | 0.1466 | 0.198
|  |  |  | 30 | 1:3 | W | Wash | 0.0013 | 0 | 0.002
| 2019-09-06 | 1 | a | 31 | 1:25 | 1 | Unknown | 0.0086 | 0.0074 | 0.012
| 2019-09-06 | 1 | b | 32 | 1:26 | 2 | Unknown | 0.0083 | 0.0073 | 0.012
| 2019-09-06 | 1 | c | 33 | 1:27 | 3 | Unknown | 0.0078 | 0.0069 | 0.011
| 2019-09-06 | 2 | a | 34 | 1:28 | 4 | Unknown | 0.0076 | 0.0069 | 0.011
| 2019-09-06 | 2 | b | 35 | 1:29 | 5 | Unknown | 0.0073 | 0.0067 | 0.011
| 2019-09-06 | 2 | c | 37 | 1:30 | 6 | Unknown | 0.006 | 0.0057 | 0.01
| 2019-09-06 | 3 | a | 38 | 1:31 | 7 | Unknown | 0.0291 | 0.0287 | 0.04
| 2019-09-06 | 3 | b | 39 | 1:32 | 8 | Unknown | 0.026 | 0.0254 | 0.036
| 2019-09-06 | 3 | c | 40 | 1:33 | 9 | Unknown | 0.025 | 0.0243 | 0.035
| 2019-09-06 | 4 | a | 41 | 1:34 | 10 | Unknown | 0.0044 | 0.0038 | 0.007
| 2019-09-06 | 4 | b | 42 | 1:35 | 11 | Unknown | 0.004 | 0.0035 | 0.007
| 2019-09-06 | 4 | c | 43 | 1:36 | 12 | Unknown | 0.0041 | 0.0036 | 0.007
| 2019-09-06 | 5 | a | 44 | 1:37 | 13 | Unknown | 0.0375 | 0.037 | 0.052
| 2019-09-06 | 5 | b | 45 | 1:38 | 14 | Unknown | 0.0311 | 0.0302 | 0.043
| 2019-09-06 | 5 | c | 46 | 1:39 | 15 | Unknown | 0.0369 | 0.0361 | 0.05
| 2019-09-06 | 6 | a | 50 | 1:40 | 16 | Unknown | 0.1149 | 0.1137 | 0.154
| 2019-09-06 | 6 | b | 51 | 1:41 | 17 | Unknown | 0.123 | 0.1216 | 0.165
| 2019-09-06 | 6 | c | 52 | 1:42 | 18 | Unknown | 0.0969 | 0.0954 | 0.13
| 2019-09-06 | 7 | a | 53 | 1:43 | 19 | Unknown | 0.0789 | 0.0776 | 0.106
| 2019-09-06 | 7 | b | 54 | 1:44 | 20 | Unknown | 0.0657 | 0.0644 | 0.088
| 2019-09-06 | 7 | c | 55 | 1:45 | 21 | Unknown | 0.0611 | 0.06 | 0.082
| 2019-09-06 | 8 | a | 56 | 1:46 | 22 | Unknown | 0.0413 | 0.0402 | 0.056
| 2019-09-06 | 8 | b | 57 | 1:47 | 23 | Unknown | 0.0355 | 0.0345 | 0.048
| 2019-09-06 | 8 | c | 59 | 1:48 | 24 | Unknown | 0.0258 | 0.0251 | 0.036
| 2019-09-06 | 9 | a | 60 | 1:49 | 25 | Unknown | 0.0051 | 0.0042 | 0.008
| 2019-09-06 | 9 | b | 61 | 1:50 | 26 | Unknown | 0.0043 | 0.0036 | 0.007
| 2019-09-06 | 9 | c | 62 | 1:51 | 27 | Unknown | 0.005 | 0.0044 | 0.008
| 2019-09-06 | 10 | a | 63 | 1:52 | 28 | Unknown | 0.004 | 0.0034 | 0.007
| 2019-09-06 | 10 | b | 64 | 1:53 | 29 | Unknown | 0.0047 | 0.0041 | 0.008
| 2019-09-06 | 10 | c | 67 | 1:54 | 30 | Unknown | 0.0051 | 0.0041 | 0.008
| 2019-09-06 | 11 | a | 68 | 1:55 | 31 | Unknown | 0.0028 | 0.0024 | 0.005
| 2019-09-06 | 11 | b | 70 | 1:56 | 32 | Unknown | 0.0029 | 0.0024 | 0.005
| 2019-09-06 | 11 | c | 71 | 1:57 | 33 | Unknown | 0.0026 | 0.0021 | 0.005
| 2019-09-06 | 12 | a | 72 | 1:58 | 34 | Unknown | 0.0045 | 0.0041 | 0.008
| 2019-09-06 | 12 | b | 73 | 1:59 | 35 | Unknown | 0.0047 | 0.0042 | 0.008
| 2019-09-06 | 12 | c | 74 | 1:60 | 36 | Unknown | 0.0047 | 0.0042 | 0.008
| 2019-09-06 | 13 | a | 75 | 2:1 | 37 | Unknown | 0.0025 | 0.002 | 0.005
| 2019-09-06 | 13 | b | 76 | 2:2 | 38 | Unknown | 0.0018 | 0.0013 | 0.004
| 2019-09-06 | 13 | c | 77 | 2:3 | 39 | Unknown | 0.0021 | 0.0016 | 0.004
| 2019-09-17 | 1 | a | 78 | 2:4 | 40 | Unknown | 0.0113 | 0.0108 | 0.017
| 2019-09-17 | 1 | b | 79 | 2:5 | 41 | Unknown | 0.0115 | 0.0109 | 0.017
| 2019-09-17 | 1 | c | 81 | 2:6 | 42 | Unknown | 0.0101 | 0.0096 | 0.015
| 2019-09-17 | 2 | a | 82 | 2:7 | 43 | Unknown | 0.0063 | 0.0057 | 0.01
| 2019-09-17 | 2 | b | 83 | 2:8 | 44 | Unknown | 0.0057 | 0.0051 | 0.009
| 2019-09-17 | 2 | c | 84 | 2:9 | 45 | Unknown | 0.0068 | 0.0062 | 0.011
| 2019-09-17 | 3 | a | 87 | 2:10 | 46 | Unknown | 0.0231 | 0.022 | 0.032
| 2019-09-17 | 3 | b | 88 | 2:11 | 47 | Unknown | 0.028 | 0.0273 | 0.039
| 2019-09-17 | 3 | c | 89 | 2:12 | 48 | Unknown | 0.0409 | 0.0401 | 0.056
| 2019-09-17 | 4 | a | 90 | 2:13 | 49 | Unknown | 0.0047 | 0.0038 | 0.007
| 2019-09-17 | 4 | b | 92 | 2:14 | 50 | Unknown | 0.0045 | 0.0039 | 0.007
| 2019-09-17 | 4 | c | 95 | 2:15 | 51 | Unknown | 0.0055 | 0.0037 | 0.007
| 2019-09-17 | 5 | a | 96 | 2:16 | 52 | Unknown | 0.0345 | 0.0337 | 0.047
| 2019-09-17 | 5 | b | 97 | 2:17 | 53 | Unknown | 0.0339 | 0.0328 | 0.046
| 2019-09-17 | 5 | c | 98 | 2:18 | 54 | Unknown | 0.0243 | 0.0231 | 0.033
| 2019-09-17 | 6 | a | 99 | 2:19 | 55 | Unknown | 0.1054 | 0.1043 | 0.141
| 2019-09-17 | 6 | b | 100 | 2:20 | 56 | Unknown | 0.1068 | 0.1051 | 0.143
| 2019-09-17 | 6 | c | 101 | 2:21 | 57 | Unknown | 0.1019 | 0.1001 | 0.136
| 2019-09-17 | 7 | a | 103 | 2:22 | 58 | Unknown | 0.1137 | 0.1127 | 0.153
| 2019-09-17 | 7 | b | 104 | 2:23 | 59 | Unknown | 0.097 | 0.0952 | 0.129
| 2019-09-17 | 7 | c | 105 | 2:24 | 60 | Unknown | 0.0836 | 0.082 | 0.112
| 2019-09-17 | 8 | a | 108 | 2:25 | 61 | Unknown | 0.0904 | 0.0891 | 0.121
| 2019-09-17 | 8 | b | 109 | 2:26 | 62 | Unknown | 0.0717 | 0.0704 | 0.096
| 2019-09-17 | 8 | c | 110 | 2:27 | 63 | Unknown | 0.0621 | 0.0611 | 0.084
| 2019-09-17 | 9 | a | 111 | 2:28 | 64 | Unknown | 0.0154 | 0.0144 | 0.021
| 2019-09-17 | 9 | b | 112 | 2:29 | 65 | Unknown | 0.0162 | 0.0156 | 0.023
| 2019-09-17 | 9 | c | 114 | 2:30 | 66 | Unknown | 0.0134 | 0.013 | 0.02
| 2019-09-17 | 10 | a | 115 | 2:31 | 67 | Unknown | 0.0101 | 0.0096 | 0.015
| 2019-09-17 | 10 | b | 116 | 2:32 | 68 | Unknown | 0.0077 | 0.0073 | 0.012
| 2019-09-17 | 10 | c | 117 | 2:33 | 69 | Unknown | 0.0091 | 0.0088 | 0.014
| 2019-09-17 | 11 | a | 118 | 2:34 | 70 | Unknown | 0.0048 | 0.0045 | 0.008
| 2019-09-17 | 11 | b | 119 | 2:35 | 71 | Unknown | 0.0043 | 0.004 | 0.007
| 2019-09-17 | 11 | c | 120 | 2:36 | 72 | Unknown | 0.0033 | 0.003 | 0.006
| 2019-09-17 | 12 | a | 121 | 2:37 | 73 | Unknown | 0.0055 | 0.0052 | 0.009
| 2019-09-17 | 12 | b | 122 | 2:38 | 74 | Unknown | 0.0061 | 0.0058 | 0.01
| 2019-09-17 | 12 | c | 123 | 2:39 | 75 | Unknown | 0.0058 | 0.0055 | 0.01
| 2019-09-17 | 13 | a | 127 | 2:40 | 76 | Unknown | 0.0021 | 0.0011 | 0.004
| 2019-09-17 | 13 | b | 128 | 2:41 | 77 | Unknown | 0.0016 | 0.0011 | 0.004
| 2019-09-17 | 13 | c | 129 | 2:42 | 78 | Unknown | 0.0015 | 0.0009 | 0.003
| 2019-10-04 | 1 | a | 130 | 2:43 | 79 | Unknown | 0.0173 | 0.0167 | 0.025
| 2019-10-04 | 1 | b | 131 | 2:44 | 80 | Unknown | 0.0189 | 0.0181 | 0.026
| 2019-10-04 | 1 | c | 132 | 2:45 | 81 | Unknown | 0.0155 | 0.0146 | 0.022
| 2019-10-04 | 2 | a | 136 | 2:46 | 82 | Unknown | 0.0097 | 0.0087 | 0.014
| 2019-10-04 | 2 | b | 137 | 2:47 | 83 | Unknown | 0.0084 | 0.0073 | 0.012
| 2019-10-04 | 2 | c | 138 | 2:48 | 84 | Unknown | 0.0098 | 0.0088 | 0.014
| 2019-10-04 | 3 | a | 139 | 2:49 | 85 | Unknown | 0.0398 | 0.0387 | 0.054
| 2019-10-04 | 3 | b | 140 | 2:50 | 86 | Unknown | 0.0394 | 0.0381 | 0.053
| 2019-10-04 | 3 | c | 141 | 2:51 | 87 | Unknown | 0.0446 | 0.0433 | 0.06
| 2019-10-04 | 4 | a | 142 | 2:52 | 88 | Unknown | 0.0135 | 0.0122 | 0.018
| 2019-10-04 | 4 | b | 143 | 2:53 | 89 | Unknown | 0.0127 | 0.0116 | 0.018
| 2019-10-04 | 4 | c | 147 | 2:54 | 90 | Unknown | 0.0151 | 0.014 | 0.021
| 2019-10-04 | 5 | a | 148 | 2:55 | 91 | Unknown | 0.0314 | 0.0303 | 0.043
| 2019-10-04 | 5 | b | 149 | 2:56 | 92 | Unknown | 0.0309 | 0.0297 | 0.042
| 2019-10-04 | 5 | c | 150 | 2:57 | 93 | Unknown | 0.0286 | 0.0274 | 0.039
| 2019-10-04 | 6 | a | 151 | 2:58 | 94 | Unknown | 0.0807 | 0.0796 | 0.108
| 2019-10-04 | 6 | b | 152 | 2:59 | 95 | Unknown | 0.0601 | 0.0587 | 0.081
| 2019-10-04 | 6 | c | 153 | 2:60 | 96 | Unknown | 0.0684 | 0.0671 | 0.092
| 2019-10-04 | 7 | a | 154 | 3:1 | 97 | Unknown | 0.0618 | 0.0605 | 0.083
| 2019-10-04 | 7 | b | 155 | 3:2 | 98 | Unknown | 0.0476 | 0.0463 | 0.064
| 2019-10-04 | 7 | c | 156 | 3:3 | 99 | Unknown | 0.046 | 0.0449 | 0.062
| 2019-10-04 | 8 | a | 158 | 3:4 | 100 | Unknown | 0.039 | 0.0383 | 0.053
| 2019-10-04 | 8 | b | 159 | 3:5 | 101 | Unknown | 0.0454 | 0.0445 | 0.062
| 2019-10-04 | 8 | c | 160 | 3:6 | 102 | Unknown | 0.0463 | 0.0453 | 0.063
| 2019-10-04 | 9 | a | 161 | 3:7 | 103 | Unknown | 0.003 | 0.002 | 0.005
| 2019-10-04 | 9 | b | 162 | 3:8 | 104 | Unknown | 0.0028 | 0.0021 | 0.005
| 2019-10-04 | 9 | c | 163 | 3:9 | 105 | Unknown | 0.0031 | 0.0025 | 0.006
| 2019-10-04 | 10 | a | 164 | 3:10 | 106 | Unknown | 0.0035 | 0.0029 | 0.006
| 2019-10-04 | 10 | b | 165 | 3:11 | 107 | Unknown | 0.0035 | 0.003 | 0.006
| 2019-10-04 | 10 | c | 166 | 3:12 | 108 | Unknown | 0.0034 | 0.0028 | 0.006
| 2019-10-04 | 11 | a | 167 | 3:13 | 109 | Unknown | 0.0026 | 0.0021 | 0.005
| 2019-10-04 | 11 | b | 171 | 3:14 | 110 | Unknown | 0.0031 | 0.002 | 0.005
| 2019-10-04 | 11 | c | 172 | 3:15 | 111 | Unknown | 0.0026 | 0.0021 | 0.005
| 2019-10-04 | 12 | a | 173 | 3:16 | 112 | Unknown | 0.0046 | 0.0041 | 0.008
| 2019-10-04 | 12 | b | 174 | 3:17 | 113 | Unknown | 0.0035 | 0.003 | 0.006
| 2019-10-04 | 12 | c | 175 | 3:18 | 114 | Unknown | 0.0038 | 0.0034 | 0.007
| 2019-10-04 | 13 | a | 176 | 3:19 | 115 | Unknown | 0.002 | 0.0015 | 0.004
| 2019-10-04 | 13 | b | 177 | 3:20 | 116 | Unknown | 0.0023 | 0.0018 | 0.005
| 2019-10-04 | 13 | c | 178 | 3:21 | 117 | Unknown | 0.0021 | 0.0016 | 0.004
| 2019-09-06 | 8a | a | 182 | 3:22 | 118 | Unknown | 0.0047 | 0.0032 | 0.006
| 2019-09-06 | 8a | b | 183 | 3:23 | 119 | Unknown | 0.0044 | 0.0039 | 0.007
| 2019-09-06 | 8a | c | 184 | 3:24 | 120 | Unknown | 0.0355 | 0.035 | 0.049
| 2019-09-06 | 9a | a | 185 | 3:25 | 121 | Unknown | 0.0479 | 0.0472 | 0.065
| 2019-09-06 | 9a | b | 186 | 3:26 | 122 | Unknown | 0.0026 | 0.0018 | 0.005
| 2019-09-06 | 9a | c | 187 | 3:27 | 123 | Unknown | 0.0308 | 0.0303 | 0.043
| 2019-09-17 | 8a | a | 188 | 3:28 | 124 | Unknown | 0.0016 | 0.0009 | 0.003
| 2019-09-17 | 8a | b | 189 | 3:29 | 125 | Unknown | 0.0014 | 0.0009 | 0.003
| 2019-09-17 | 8a | c | 191 | 3:30 | 126 | Unknown | 0.0019 | 0.0014 | 0.004
| 2019-09-17 | 9a | a | 192 | 3:31 | 127 | Unknown | 0.0026 | 0.0021 | 0.005
| 2019-09-17 | 9a | b | 195 | 3:32 | 128 | Unknown | 0.0044 | 0.0033 | 0.007
| 2019-09-17 | 9a | c | 196 | 3:33 | 129 | Unknown | 0.0037 | 0.0031 | 0.006
| 2019-10-04 | 8a | a | 197 | 3:34 | 130 | Unknown | 0.0031 | 0.0025 | 0.005
| 2019-10-04 | 8a | b | 198 | 3:35 | 131 | Unknown | 0.0028 | 0.0022 | 0.005
| 2019-10-04 | 8a | c | 199 | 3:36 | 132 | Unknown | 0.0029 | 0.0022 | 0.005
| 2019-10-04 | 9a | a | 200 | 3:37 | 133 | Unknown | 0.0023 | 0.0017 | 0.004
| 2019-10-04 | 9a | b | 204 | 3:38 | 134 | Unknown | 0.0031 | 0.0015 | 0.004
| 2019-10-04 | 9a | c | 205 | 3:39 | 135 | Unknown | 0.0022 | 0.0016 | 0.004
| 2019-10-04 | Blank | a | 206 | 3:40 | 136 | Unknown | 0.0022 | 0.0016 | 0.004

## Std Curve Construction
1. Copy in Calibrant Information
[10 empty] | 2. Run Linear Regression | [8 empty] | 4. If quadratic exceeds 10%, run split-curve procedure.
Use X=Peak height; Y=Conc)
| Calibrant |  | Linear Fit |  | Quadratic |  | Split Curve
Calibrant | Conc. | Calibrant | Std Conc |  | Std Conc |  | Std Conc | [11 empty] | Plot low range calibrant data: |  |  | Enter coefficients below graph.
Id | (mg/L N) | peak hght (x) | (y-hat) | % error | (y-hat) | % error | (y-hat) | % error |  |  | All data - Linear Fit
C1 | 0.001 | 0.0002 | 0.0024478 | 144.78 | 0.0017746 | 77.4636 | 0.001316 | 31.6017 |  | SUMMARY OUTPUT
C1 | 0.001 | 0 | 0.0021807 | 118.0722 | 0.0015 | 50 | 0.001 | 0
C1 | 0.001 | -0.0001 | 0.0020472 | 104.7184 | 0.0013627 | 36.2679 | 0.00084182 | -15.8176 |  | Regression Statistics
C2 | 0.005 | 0.0026 | 0.0056527 | 13.0546 | 0.0050697 | 1.393 | 0.0050734 | 1.4679 |  | Multiple R | 0.99998
C2 | 0.005 | 0.0024 | 0.0053857 | 7.7131 | 0.0047951 | -4.0978 | 0.0047627 | -4.7453 |  | R Square | 0.99997
C2 | 0.005 | 0.0024 | 0.0053857 | 7.7131 | 0.0047951 | -4.0978 | 0.0047627 | -4.7453 |  | Adjusted R Square | 0.99996
C3 | 0.01 | 0.006 | 0.010193 | 1.9305 | 0.0097356 | -2.6436 | 0.010286 | 2.8629 |  | Standard Error | 0.0010987
C3 | 0.01 | 0.0058 | 0.009926 | -0.74028 | 0.0094612 | -5.3877 | 0.0099832 | -0.16775 |  | Observations | 18
C3 | 0.01 | 0.0057 | 0.0097924 | -2.0757 | 0.009324 | -6.7597 | 0.0098315 | -1.6848
C4 | 0.05 | 0.0352 | 0.049186 | -1.6272 | 0.049714 | -0.5718 | 0.049744 | -0.51294 |  | ANOVA
C4 | 0.05 | 0.0354 | 0.049453 | -1.0931 | 0.049987 | -0.025315 | 0.049981 | -0.03806 |  |  | df | SS | MS | F | Significance F
C4 | 0.05 | 0.0356 | 0.049721 | -0.55894 | 0.050261 | 0.52116 | 0.050218 | 0.43593 | Split | Regression | 1 | 0.56608 | 0.56608 | 468954.1757 | 3.605e-37
C5 | 0.1 | 0.0721 | 0.098462 | -1.5378 | 0.099994 | -0.0064027 | 0.099952 | -0.047516 |  | Residual | 16 | 1.9314e-05 | 1.2071e-06
C5 | 0.1 | 0.0723 | 0.098729 | -1.2707 | 0.10027 | 0.26538 | 0.10022 | 0.22418 |  | Total | 17 | 0.5661
C5 | 0.1 | 0.0721 | 0.098462 | -1.5378 | 0.099994 | -0.0064027 | 0.099952 | -0.047516
C6 | 0.5 | 0.3733 | 0.50068 | 0.13621 | 0.50033 | 0.066712 | 0.50032 | 0.064617 |  |  | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95.0% | Upper 95.0%
C6 | 0.5 | 0.3742 | 0.50188 | 0.37658 | 0.5015 | 0.30058 | 0.50149 | 0.2986 |  | Intercept | 0.0021807 | 0.00030383 | 7.1774 | 2.1996e-06 | 0.0015366 | 0.0028248 | 0.0015366 | 0.0028248 | Low curve coefficients:
C6 | 0.5 | 0.3716 | 0.49841 | -0.31783 | 0.49812 | -0.37513 | 0.49811 | -0.37744 |  | X Variable 1 | 1.3354 | 0.00195 | 684.8023 | 3.605e-37 | 1.3313 | 1.3395 | 1.3313 | 1.3395 |  | intcpt | X coef. | X^2 coef.
[20 empty] | 0.001 | 1.5812 | -5.5807
[10 empty] | (Use X=Peak Ht.; Y=Conc) | [8 empty] | Plot High Range calibrant data. |  |  | Enter coefficients below graph.
Enter equation parameters below graph.
[10 empty] | Quadratic curve coefficients | [8 empty] | High curve coefficients:
[13 empty] | Intercept | X coef. | X^2 coef. |  |  |  |  | intcpt | X coef | X^2 coef
[13 empty] | 0.0015 | 1.3732 | -0.0989 |  |  |  |  | 0.0015 | 1.3725 | -0.0971
Show "Split" in spreadsheet.

## Apply Std Curve
|  |  | Sample Info |  |  |  | FASPacN-NO3 |  |  | Curve-fit results N-NO3
SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | Peak# | Position | Identifier | Type | Raw Ht | Cor Ht | mg/L | mg/L
2019-09-06 | 1 | a | 31 | 1:25 | 1 | Unknown | 0.0086 | 0.0074 | 0.012 | 0.012395 |  | Curve-fitting parameters used:
2019-09-06 | 1 | b | 32 | 1:26 | 2 | Unknown | 0.0083 | 0.0073 | 0.012 | 0.012245 |  | Copy from Std Curve Construction page.
2019-09-06 | 1 | c | 33 | 1:27 | 3 | Unknown | 0.0078 | 0.0069 | 0.011 | 0.011645
2019-09-06 | 2 | a | 34 | 1:28 | 4 | Unknown | 0.0076 | 0.0069 | 0.011 | 0.011645 |  | Use these for linear or quadratic fit.
2019-09-06 | 2 | b | 35 | 1:29 | 5 | Unknown | 0.0073 | 0.0067 | 0.011 | 0.011344 |  | If linear, enter 0 for X^2 coefficient.
2019-09-06 | 2 | c | 37 | 1:30 | 6 | Unknown | 0.006 | 0.0057 | 0.01 | 0.0098315
2019-09-06 | 3 | a | 38 | 1:31 | 7 | Unknown | 0.0291 | 0.0287 | 0.04 | 0.041784 |  | Intercept: | X Coef. | X^2 Coef:
2019-09-06 | 3 | b | 39 | 1:32 | 8 | Unknown | 0.026 | 0.0254 | 0.036 | 0.037562 |  | 5.9731e-05 | 1.3228
2019-09-06 | 3 | c | 40 | 1:33 | 9 | Unknown | 0.025 | 0.0243 | 0.035 | 0.036128 |  | -0.0001 | 1.3335 | -0.0275
2019-09-06 | 4 | a | 41 | 1:34 | 10 | Unknown | 0.0044 | 0.0038 | 0.007 | 0.006928
2019-09-06 | 4 | b | 42 | 1:35 | 11 | Unknown | 0.004 | 0.0035 | 0.007 | 0.0064658 |  | Use these for split curve:
2019-09-06 | 4 | c | 43 | 1:36 | 12 | Unknown | 0.0041 | 0.0036 | 0.007 | 0.00662 |  | Low Range:
2019-09-06 | 5 | a | 44 | 1:37 | 13 | Unknown | 0.0375 | 0.037 | 0.052 | 0.05215 |  | Intercept: | X Coef. | X^2 Coef:
2019-09-06 | 5 | b | 45 | 1:38 | 14 | Unknown | 0.0311 | 0.0302 | 0.043 | 0.043662 |  | 0.001 | 1.5812 | -5.5807
2019-09-06 | 5 | c | 46 | 1:39 | 15 | Unknown | 0.0369 | 0.0361 | 0.05 | 0.050808
2019-09-06 | 6 | a | 50 | 1:40 | 16 | Unknown | 0.1149 | 0.1137 | 0.154 | 0.1563
2019-09-06 | 6 | b | 51 | 1:41 | 17 | Unknown | 0.123 | 0.1216 | 0.165 | 0.16696 |  | High Range:
2019-09-06 | 6 | c | 52 | 1:42 | 18 | Unknown | 0.0969 | 0.0954 | 0.13 | 0.13155 |  | Intercept: | X Coef. | X^2 Coef:
2019-09-06 | 7 | a | 53 | 1:43 | 19 | Unknown | 0.0789 | 0.0776 | 0.106 | 0.10742 |  | 0.0015 | 1.3725 | -0.0971
2019-09-06 | 7 | b | 54 | 1:44 | 20 | Unknown | 0.0657 | 0.0644 | 0.088 | 0.089486
2019-09-06 | 7 | c | 55 | 1:45 | 21 | Unknown | 0.0611 | 0.06 | 0.082 | 0.0835 |  | Procedure for split curve:
2019-09-06 | 8 | a | 56 | 1:46 | 22 | Unknown | 0.0413 | 0.0402 | 0.056 | 0.056518 |  | 1. Sort data on Row I: FASpac results, low to high.
2019-09-06 | 8 | b | 57 | 1:47 | 23 | Unknown | 0.0355 | 0.0345 | 0.048 | 0.048909 |  | 2. Apply low range equation.
2019-09-06 | 8 | c | 59 | 1:48 | 24 | Unknown | 0.0258 | 0.0251 | 0.036 | 0.037172 |  | 3. Apply high range equation.
2019-09-06 | 9 | a | 60 | 1:49 | 25 | Unknown | 0.0051 | 0.0042 | 0.008 | 0.0075426 |  | 4. Note here peak height where split occurs: |  |  | 0.05
2019-09-06 | 9 | b | 61 | 1:50 | 26 | Unknown | 0.0043 | 0.0036 | 0.007 | 0.00662 |  | 4. Sort back based on Row D: Peak Number
2019-09-06 | 9 | c | 62 | 1:51 | 27 | Unknown | 0.005 | 0.0044 | 0.008 | 0.0078492
2019-09-06 | 10 | a | 63 | 1:52 | 28 | Unknown | 0.004 | 0.0034 | 0.007 | 0.0063116
2019-09-06 | 10 | b | 64 | 1:53 | 29 | Unknown | 0.0047 | 0.0041 | 0.008 | 0.0073891
2019-09-06 | 10 | c | 67 | 1:54 | 30 | Unknown | 0.0051 | 0.0041 | 0.008 | 0.0073891
2019-09-06 | 11 | a | 68 | 1:55 | 31 | Unknown | 0.0028 | 0.0024 | 0.005 | 0.0047627
2019-09-06 | 11 | b | 70 | 1:56 | 32 | Unknown | 0.0029 | 0.0024 | 0.005 | 0.0047627
2019-09-06 | 11 | c | 71 | 1:57 | 33 | Unknown | 0.0026 | 0.0021 | 0.005 | 0.0042959
2019-09-06 | 12 | a | 72 | 1:58 | 34 | Unknown | 0.0045 | 0.0041 | 0.008 | 0.0073891
2019-09-06 | 12 | b | 73 | 1:59 | 35 | Unknown | 0.0047 | 0.0042 | 0.008 | 0.0075426
2019-09-06 | 12 | c | 74 | 1:60 | 36 | Unknown | 0.0047 | 0.0042 | 0.008 | 0.0075426
2019-09-06 | 13 | a | 75 | 2:1 | 37 | Unknown | 0.0025 | 0.002 | 0.005 | 0.0041401
2019-09-06 | 13 | b | 76 | 2:2 | 38 | Unknown | 0.0018 | 0.0013 | 0.004 | 0.0030461
2019-09-06 | 13 | c | 77 | 2:3 | 39 | Unknown | 0.0021 | 0.0016 | 0.004 | 0.0035156
2019-09-17 | 1 | a | 78 | 2:4 | 40 | Unknown | 0.0113 | 0.0108 | 0.017 | 0.017426
2019-09-17 | 1 | b | 79 | 2:5 | 41 | Unknown | 0.0115 | 0.0109 | 0.017 | 0.017572
2019-09-17 | 1 | c | 81 | 2:6 | 42 | Unknown | 0.0101 | 0.0096 | 0.015 | 0.015665
2019-09-17 | 2 | a | 82 | 2:7 | 43 | Unknown | 0.0063 | 0.0057 | 0.01 | 0.0098315
2019-09-17 | 2 | b | 83 | 2:8 | 44 | Unknown | 0.0057 | 0.0051 | 0.009 | 0.008919
2019-09-17 | 2 | c | 84 | 2:9 | 45 | Unknown | 0.0068 | 0.0062 | 0.011 | 0.010589
2019-09-17 | 3 | a | 87 | 2:10 | 46 | Unknown | 0.0231 | 0.022 | 0.032 | 0.033085
2019-09-17 | 3 | b | 88 | 2:11 | 47 | Unknown | 0.028 | 0.0273 | 0.039 | 0.040008
2019-09-17 | 3 | c | 89 | 2:12 | 48 | Unknown | 0.0409 | 0.0401 | 0.056 | 0.056381
2019-09-17 | 4 | a | 90 | 2:13 | 49 | Unknown | 0.0047 | 0.0038 | 0.007 | 0.006928
2019-09-17 | 4 | b | 92 | 2:14 | 50 | Unknown | 0.0045 | 0.0039 | 0.007 | 0.0070818
2019-09-17 | 4 | c | 95 | 2:15 | 51 | Unknown | 0.0055 | 0.0037 | 0.007 | 0.006774
2019-09-17 | 5 | a | 96 | 2:16 | 52 | Unknown | 0.0345 | 0.0337 | 0.047 | 0.047948
2019-09-17 | 5 | b | 97 | 2:17 | 53 | Unknown | 0.0339 | 0.0328 | 0.046 | 0.046859
2019-09-17 | 5 | c | 98 | 2:18 | 54 | Unknown | 0.0243 | 0.0231 | 0.033 | 0.034548
2019-09-17 | 6 | a | 99 | 2:19 | 55 | Unknown | 0.1054 | 0.1043 | 0.141 | 0.1436
2019-09-17 | 6 | b | 100 | 2:20 | 56 | Unknown | 0.1068 | 0.1051 | 0.143 | 0.14468
2019-09-17 | 6 | c | 101 | 2:21 | 57 | Unknown | 0.1019 | 0.1001 | 0.136 | 0.13791
2019-09-17 | 7 | a | 103 | 2:22 | 58 | Unknown | 0.1137 | 0.1127 | 0.153 | 0.15495
2019-09-17 | 7 | b | 104 | 2:23 | 59 | Unknown | 0.097 | 0.0952 | 0.129 | 0.13128
2019-09-17 | 7 | c | 105 | 2:24 | 60 | Unknown | 0.0836 | 0.082 | 0.112 | 0.11339
2019-09-17 | 8 | a | 108 | 2:25 | 61 | Unknown | 0.0904 | 0.0891 | 0.121 | 0.12302
2019-09-17 | 8 | b | 109 | 2:26 | 62 | Unknown | 0.0717 | 0.0704 | 0.096 | 0.097643
2019-09-17 | 8 | c | 110 | 2:27 | 63 | Unknown | 0.0621 | 0.0611 | 0.084 | 0.084997
2019-09-17 | 9 | a | 111 | 2:28 | 64 | Unknown | 0.0154 | 0.0144 | 0.021 | 0.022612
2019-09-17 | 9 | b | 112 | 2:29 | 65 | Unknown | 0.0162 | 0.0156 | 0.023 | 0.024309
2019-09-17 | 9 | c | 114 | 2:30 | 66 | Unknown | 0.0134 | 0.013 | 0.02 | 0.020612
2019-09-17 | 10 | a | 115 | 2:31 | 67 | Unknown | 0.0101 | 0.0096 | 0.015 | 0.015665
2019-09-17 | 10 | b | 116 | 2:32 | 68 | Unknown | 0.0077 | 0.0073 | 0.012 | 0.012245
2019-09-17 | 10 | c | 117 | 2:33 | 69 | Unknown | 0.0091 | 0.0088 | 0.014 | 0.014482
2019-09-17 | 11 | a | 118 | 2:34 | 70 | Unknown | 0.0048 | 0.0045 | 0.008 | 0.0080024
2019-09-17 | 11 | b | 119 | 2:35 | 71 | Unknown | 0.0043 | 0.004 | 0.007 | 0.0072355
2019-09-17 | 11 | c | 120 | 2:36 | 72 | Unknown | 0.0033 | 0.003 | 0.006 | 0.0056934
2019-09-17 | 12 | a | 121 | 2:37 | 73 | Unknown | 0.0055 | 0.0052 | 0.009 | 0.0090713
2019-09-17 | 12 | b | 122 | 2:38 | 74 | Unknown | 0.0061 | 0.0058 | 0.01 | 0.0099832
2019-09-17 | 12 | c | 123 | 2:39 | 75 | Unknown | 0.0058 | 0.0055 | 0.01 | 0.0095278
2019-09-17 | 13 | a | 127 | 2:40 | 76 | Unknown | 0.0021 | 0.0011 | 0.004 | 0.0027326
2019-09-17 | 13 | b | 128 | 2:41 | 77 | Unknown | 0.0016 | 0.0011 | 0.004 | 0.0027326
2019-09-17 | 13 | c | 129 | 2:42 | 78 | Unknown | 0.0015 | 0.0009 | 0.003 | 0.0024186
2019-10-04 | 1 | a | 130 | 2:43 | 79 | Unknown | 0.0173 | 0.0167 | 0.025 | 0.02585
2019-10-04 | 1 | b | 131 | 2:44 | 80 | Unknown | 0.0189 | 0.0181 | 0.026 | 0.027791
2019-10-04 | 1 | c | 132 | 2:45 | 81 | Unknown | 0.0155 | 0.0146 | 0.022 | 0.022896
2019-10-04 | 2 | a | 136 | 2:46 | 82 | Unknown | 0.0097 | 0.0087 | 0.014 | 0.014334
2019-10-04 | 2 | b | 137 | 2:47 | 83 | Unknown | 0.0084 | 0.0073 | 0.012 | 0.012245
2019-10-04 | 2 | c | 138 | 2:48 | 84 | Unknown | 0.0098 | 0.0088 | 0.014 | 0.014482
2019-10-04 | 3 | a | 139 | 2:49 | 85 | Unknown | 0.0398 | 0.0387 | 0.054 | 0.05447
2019-10-04 | 3 | b | 140 | 2:50 | 86 | Unknown | 0.0394 | 0.0381 | 0.053 | 0.053651
2019-10-04 | 3 | c | 141 | 2:51 | 87 | Unknown | 0.0446 | 0.0433 | 0.06 | 0.060747
2019-10-04 | 4 | a | 142 | 2:52 | 88 | Unknown | 0.0135 | 0.0122 | 0.018 | 0.01946
2019-10-04 | 4 | b | 143 | 2:53 | 89 | Unknown | 0.0127 | 0.0116 | 0.018 | 0.018591
2019-10-04 | 4 | c | 147 | 2:54 | 90 | Unknown | 0.0151 | 0.014 | 0.021 | 0.022043
2019-10-04 | 5 | a | 148 | 2:55 | 91 | Unknown | 0.0314 | 0.0303 | 0.043 | 0.043787
2019-10-04 | 5 | b | 149 | 2:56 | 92 | Unknown | 0.0309 | 0.0297 | 0.042 | 0.043039
2019-10-04 | 5 | c | 150 | 2:57 | 93 | Unknown | 0.0286 | 0.0274 | 0.039 | 0.040135
2019-10-04 | 6 | a | 151 | 2:58 | 94 | Unknown | 0.0807 | 0.0796 | 0.108 | 0.11014
2019-10-04 | 6 | b | 152 | 2:59 | 95 | Unknown | 0.0601 | 0.0587 | 0.081 | 0.081731
2019-10-04 | 6 | c | 153 | 2:60 | 96 | Unknown | 0.0684 | 0.0671 | 0.092 | 0.093158
2019-10-04 | 7 | a | 154 | 3:1 | 97 | Unknown | 0.0618 | 0.0605 | 0.083 | 0.084181
2019-10-04 | 7 | b | 155 | 3:2 | 98 | Unknown | 0.0476 | 0.0463 | 0.064 | 0.064839
2019-10-04 | 7 | c | 156 | 3:3 | 99 | Unknown | 0.046 | 0.0449 | 0.062 | 0.062929
2019-10-04 | 8 | a | 158 | 3:4 | 100 | Unknown | 0.039 | 0.0383 | 0.053 | 0.053924
2019-10-04 | 8 | b | 159 | 3:5 | 101 | Unknown | 0.0454 | 0.0445 | 0.062 | 0.062384
2019-10-04 | 8 | c | 160 | 3:6 | 102 | Unknown | 0.0463 | 0.0453 | 0.063 | 0.063475
2019-10-04 | 9 | a | 161 | 3:7 | 103 | Unknown | 0.003 | 0.002 | 0.005 | 0.0041401
2019-10-04 | 9 | b | 162 | 3:8 | 104 | Unknown | 0.0028 | 0.0021 | 0.005 | 0.0042959
2019-10-04 | 9 | c | 163 | 3:9 | 105 | Unknown | 0.0031 | 0.0025 | 0.006 | 0.0049181 | Split
2019-10-04 | 10 | a | 164 | 3:10 | 106 | Unknown | 0.0035 | 0.0029 | 0.006 | 0.0055385
2019-10-04 | 10 | b | 165 | 3:11 | 107 | Unknown | 0.0035 | 0.003 | 0.006 | 0.0056934
2019-10-04 | 10 | c | 166 | 3:12 | 108 | Unknown | 0.0034 | 0.0028 | 0.006 | 0.0053836
2019-10-04 | 11 | a | 167 | 3:13 | 109 | Unknown | 0.0026 | 0.0021 | 0.005 | 0.0042959
2019-10-04 | 11 | b | 171 | 3:14 | 110 | Unknown | 0.0031 | 0.002 | 0.005 | 0.0041401
2019-10-04 | 11 | c | 172 | 3:15 | 111 | Unknown | 0.0026 | 0.0021 | 0.005 | 0.0042959
2019-10-04 | 12 | a | 173 | 3:16 | 112 | Unknown | 0.0046 | 0.0041 | 0.008 | 0.0073891
2019-10-04 | 12 | b | 174 | 3:17 | 113 | Unknown | 0.0035 | 0.003 | 0.006 | 0.0056934
2019-10-04 | 12 | c | 175 | 3:18 | 114 | Unknown | 0.0038 | 0.0034 | 0.007 | 0.0063116
2019-10-04 | 13 | a | 176 | 3:19 | 115 | Unknown | 0.002 | 0.0015 | 0.004 | 0.0033592
2019-10-04 | 13 | b | 177 | 3:20 | 116 | Unknown | 0.0023 | 0.0018 | 0.005 | 0.0038281
2019-10-04 | 13 | c | 178 | 3:21 | 117 | Unknown | 0.0021 | 0.0016 | 0.004 | 0.0035156
2019-09-06 | 8a | a | 182 | 3:22 | 118 | Unknown | 0.0047 | 0.0032 | 0.006 | 0.0060027
2019-09-06 | 8a | b | 183 | 3:23 | 119 | Unknown | 0.0044 | 0.0039 | 0.007 | 0.0070818
2019-09-06 | 8a | c | 184 | 3:24 | 120 | Unknown | 0.0355 | 0.035 | 0.049 | 0.049506
2019-09-06 | 9a | a | 185 | 3:25 | 121 | Unknown | 0.0479 | 0.0472 | 0.065 | 0.066066
2019-09-06 | 9a | b | 186 | 3:26 | 122 | Unknown | 0.0026 | 0.0018 | 0.005 | 0.0038281
2019-09-06 | 9a | c | 187 | 3:27 | 123 | Unknown | 0.0308 | 0.0303 | 0.043 | 0.043787
2019-09-17 | 8a | a | 188 | 3:28 | 124 | Unknown | 0.0016 | 0.0009 | 0.003 | 0.0024186
2019-09-17 | 8a | b | 189 | 3:29 | 125 | Unknown | 0.0014 | 0.0009 | 0.003 | 0.0024186
2019-09-17 | 8a | c | 191 | 3:30 | 126 | Unknown | 0.0019 | 0.0014 | 0.004 | 0.0032027
2019-09-17 | 9a | a | 192 | 3:31 | 127 | Unknown | 0.0026 | 0.0021 | 0.005 | 0.0042959
2019-09-17 | 9a | b | 195 | 3:32 | 128 | Unknown | 0.0044 | 0.0033 | 0.007 | 0.0061572
2019-09-17 | 9a | c | 196 | 3:33 | 129 | Unknown | 0.0037 | 0.0031 | 0.006 | 0.0058481
2019-10-04 | 8a | a | 197 | 3:34 | 130 | Unknown | 0.0031 | 0.0025 | 0.005 | 0.0049181
2019-10-04 | 8a | b | 198 | 3:35 | 131 | Unknown | 0.0028 | 0.0022 | 0.005 | 0.0044516
2019-10-04 | 8a | c | 199 | 3:36 | 132 | Unknown | 0.0029 | 0.0022 | 0.005 | 0.0044516
2019-10-04 | 9a | a | 200 | 3:37 | 133 | Unknown | 0.0023 | 0.0017 | 0.004 | 0.0036719
2019-10-04 | 9a | b | 204 | 3:38 | 134 | Unknown | 0.0031 | 0.0015 | 0.004 | 0.0033592
2019-10-04 | 9a | c | 205 | 3:39 | 135 | Unknown | 0.0022 | 0.0016 | 0.004 | 0.0035156
2019-10-04 | Blank | a | 206 | 3:40 | 136 | Unknown | 0.0022 | 0.0016 | 0.004 | 0.0035156

## Final Results
SAMPLE DATE | SAMPLE SITE | SAMPLE REPLICATE | N-NO3 mg/L
2019-09-06 | 1 | a | 0.012395
2019-09-06 | 1 | b | 0.012245
2019-09-06 | 1 | c | 0.011645
2019-09-06 | 2 | a | 0.011645
2019-09-06 | 2 | b | 0.011344
2019-09-06 | 2 | c | 0.0098315
2019-09-06 | 3 | a | 0.041784
2019-09-06 | 3 | b | 0.037562
2019-09-06 | 3 | c | 0.036128
2019-09-06 | 4 | a | 0.006928
2019-09-06 | 4 | b | 0.0064658
2019-09-06 | 4 | c | 0.00662
2019-09-06 | 5 | a | 0.05215
2019-09-06 | 5 | b | 0.043662
2019-09-06 | 5 | c | 0.050808
2019-09-06 | 6 | a | 0.1563
2019-09-06 | 6 | b | 0.16696
2019-09-06 | 6 | c | 0.13155
2019-09-06 | 7 | a | 0.10742
2019-09-06 | 7 | b | 0.089486
2019-09-06 | 7 | c | 0.0835
2019-09-06 | 8 | a | 0.056518
2019-09-06 | 8 | b | 0.048909
2019-09-06 | 8 | c | 0.037172
2019-09-06 | 9 | a | 0.0075426
2019-09-06 | 9 | b | 0.00662
2019-09-06 | 9 | c | 0.0078492
2019-09-06 | 10 | a | 0.0063116
2019-09-06 | 10 | b | 0.0073891
2019-09-06 | 10 | c | 0.0073891
2019-09-06 | 11 | a | 0.0047627
2019-09-06 | 11 | b | 0.0047627
2019-09-06 | 11 | c | 0.0042959
2019-09-06 | 12 | a | 0.0073891
2019-09-06 | 12 | b | 0.0075426
2019-09-06 | 12 | c | 0.0075426
2019-09-06 | 13 | a | 0.0041401
2019-09-06 | 13 | b | 0.0030461
2019-09-06 | 13 | c | 0.0035156
2019-09-17 | 1 | a | 0.017426
2019-09-17 | 1 | b | 0.017572
2019-09-17 | 1 | c | 0.015665
2019-09-17 | 2 | a | 0.0098315
2019-09-17 | 2 | b | 0.008919
2019-09-17 | 2 | c | 0.010589
2019-09-17 | 3 | a | 0.033085
2019-09-17 | 3 | b | 0.040008
2019-09-17 | 3 | c | 0.056381
2019-09-17 | 4 | a | 0.006928
2019-09-17 | 4 | b | 0.0070818
2019-09-17 | 4 | c | 0.006774
2019-09-17 | 5 | a | 0.047948
2019-09-17 | 5 | b | 0.046859
2019-09-17 | 5 | c | 0.034548
2019-09-17 | 6 | a | 0.1436
2019-09-17 | 6 | b | 0.14468
2019-09-17 | 6 | c | 0.13791
2019-09-17 | 7 | a | 0.15495
2019-09-17 | 7 | b | 0.13128
2019-09-17 | 7 | c | 0.11339
2019-09-17 | 8 | a | 0.12302
2019-09-17 | 8 | b | 0.097643
2019-09-17 | 8 | c | 0.084997
2019-09-17 | 9 | a | 0.022612
2019-09-17 | 9 | b | 0.024309
2019-09-17 | 9 | c | 0.020612
2019-09-17 | 10 | a | 0.015665
2019-09-17 | 10 | b | 0.012245
2019-09-17 | 10 | c | 0.014482
2019-09-17 | 11 | a | 0.0080024
2019-09-17 | 11 | b | 0.0072355
2019-09-17 | 11 | c | 0.0056934
2019-09-17 | 12 | a | 0.0090713
2019-09-17 | 12 | b | 0.0099832
2019-09-17 | 12 | c | 0.0095278
2019-09-17 | 13 | a | 0.0027326
2019-09-17 | 13 | b | 0.0027326
2019-09-17 | 13 | c | 0.0024186
2019-10-04 | 1 | a | 0.02585
2019-10-04 | 1 | b | 0.027791
2019-10-04 | 1 | c | 0.022896
2019-10-04 | 2 | a | 0.014334
2019-10-04 | 2 | b | 0.012245
2019-10-04 | 2 | c | 0.014482
2019-10-04 | 3 | a | 0.05447
2019-10-04 | 3 | b | 0.053651
2019-10-04 | 3 | c | 0.060747
2019-10-04 | 4 | a | 0.01946
2019-10-04 | 4 | b | 0.018591
2019-10-04 | 4 | c | 0.022043
2019-10-04 | 5 | a | 0.043787
2019-10-04 | 5 | b | 0.043039
2019-10-04 | 5 | c | 0.040135
2019-10-04 | 6 | a | 0.11014
2019-10-04 | 6 | b | 0.081731
2019-10-04 | 6 | c | 0.093158
2019-10-04 | 7 | a | 0.084181
2019-10-04 | 7 | b | 0.064839
2019-10-04 | 7 | c | 0.062929
2019-10-04 | 8 | a | 0.053924
2019-10-04 | 8 | b | 0.062384
2019-10-04 | 8 | c | 0.063475
2019-10-04 | 9 | a | 0.0041401
2019-10-04 | 9 | b | 0.0042959
2019-10-04 | 9 | c | 0.0049181
2019-10-04 | 10 | a | 0.0055385
2019-10-04 | 10 | b | 0.0056934
2019-10-04 | 10 | c | 0.0053836
2019-10-04 | 11 | a | 0.0042959
2019-10-04 | 11 | b | 0.0041401
2019-10-04 | 11 | c | 0.0042959
2019-10-04 | 12 | a | 0.0073891
2019-10-04 | 12 | b | 0.0056934
2019-10-04 | 12 | c | 0.0063116
2019-10-04 | 13 | a | 0.0033592
2019-10-04 | 13 | b | 0.0038281
2019-10-04 | 13 | c | 0.0035156
2019-09-06 | 8a | a | 0.0060027
2019-09-06 | 8a | b | 0.0070818
2019-09-06 | 8a | c | 0.049506
2019-09-06 | 9a | a | 0.066066
2019-09-06 | 9a | b | 0.0038281
2019-09-06 | 9a | c | 0.043787
2019-09-17 | 8a | a | 0.0024186
2019-09-17 | 8a | b | 0.0024186
2019-09-17 | 8a | c | 0.0032027
2019-09-17 | 9a | a | 0.0042959
2019-09-17 | 9a | b | 0.0061572
2019-09-17 | 9a | c | 0.0058481
2019-10-04 | 8a | a | 0.0049181
2019-10-04 | 8a | b | 0.0044516
2019-10-04 | 8a | c | 0.0044516
2019-10-04 | 9a | a | 0.0036719
2019-10-04 | 9a | b | 0.0033592
2019-10-04 | 9a | c | 0.0035156
2019-10-04 | Blank | a | 0.0035156